EMEP/CCC-Report 6/2009

The twenty-sixth intercomparison of 
analytical methods within EMEP

Hilde Th. Uggerud and Anne-Gunn Hjellbrekke

Youden plot of SO4-S on impregnated filter

| NILU | : | EMEP/CCC-Report 6/2009 |
| --- | --- | --- |
| REFERENCE | : | O-7729 |
| DATE | : | AUGUST 2009 |

EMEP Co-operative Programme for Monitoring and Evaluation of the Long-range Transmission of Air Pollutants 
in Europe

The twenty-sixth intercomparison of 
analytical methods within EMEP

Hilde Th. Uggerud and Anne-Gunn Hjellbrekke

|  | Norwegian Institute for Air Research PO Box 100, NO-2027 Kjeller, Norway |
| --- | --- |

Contents

Page

1.	Introduction	5

2.	Organisation of the intercomparison	5

3.	Data handling	6

3.1	Data analysis	6

3.2	Youden plot	6

3.3	Estimating random and systematic errors from laboratory comparisons	7

3.3.1	Estimating random errors	7

3.3.2	Estimating systematic errors	9

4.	Results	9

4.1	Sulphur dioxide in absorbing solution (A-samples)	9

4.2	Sulphur dioxide and nitric acid on impregnated filter (B-samples)	10

4.3	Nitrogen dioxide in absorbing solution (C-samples)	10

4.4	Ammonia on impregnated filters (J-samples)	10

4.5	Precipitation (G-samples)	11

5.	Summary	11

6.	References	11

Appendix 1 Tables	15

Appendix 2 Figures	43

The twenty-sixth intercomparison of 
analytical methods within EMEP

# Introduction

41 different laboratories in European countries are performing chemical analysis of air and precipitation samples within EMEP (Co-operative Programme for Monitoring and Evaluation of Long-range Transmission of Air Pollutants in Europe). Since the measurement programme is based on individual national networks, the participating laboratories apply different sampling and analytical methods. Most of the methods used are described in the manual for sampling and chemical analysis (EMEP, 1996).

In order to improve the data comparability and to get a picture of the different laboratories’ performance, interlaboratory comparisons are organised by the Chemical Co-ordinating Centre (CCC) at the Norwegian Institute for Air Research (NILU). So far twenty-five intercomparisons have been arranged (Hanssen, 1988, 1990; Hanssen et al., 1983; Hanssen and Ladegård, 1984, 1985, 1987; Hanssen and Skjelmoen, 1992, 1994, 1995, 1996, 1997, 2001; Thrane, 1978, 1980a, 1980b, 1981; Uggerud et al., 2001, 2002, 2003, 2004; Hjellbrekke et al., 2005; Uggerud and Hjellbrekke, 2007, 2008).

Since 2000 the laboratory intercomparisons within EMEP have also been open for participation of laboratories from other networks.

- This report gives the results of the twenty-sixth interlaboratory test.

# Organisation of the intercomparison

The samples for the twenty-sixth intercomparison (see Table 2) were prepared and distributed to 91 laboratories in June 2008.

Most of the laboratories had returned their results to the CCC within one month after the deadline given as 15 November 2008. A total of 68 laboratories have returned their results. This includes 36 EMEP-laboratories.

Expected values were released at EMEP home page in Januaery 2009. In accordance with the decision of the Steering Body of EMEP, the results are presented in such a way that the different laboratories are identified. Tables 3a and 3b give the names of the participating laboratories together with the numbers used when presenting the results in tables and figures.

Information received on the analytical methods used is given in Tables 4–9.

# Data handling

The data reported from the participants are presented in Tables 13-28 and Figures 2-17. An overview of all results is presented in Tables 10 and 11.

## Data analysis

The reported values are presented in the tables in decreasing order together with the number of the laboratory. The expected (theoretical) value, the number of results, the arithmetic mean value, the median, the standard deviation and the relative standard deviation in percent are also given. After the first statistical run with all results included, the calculation was repeated with the outliers excluded. The outliers (unused) are defined as the results more than two standard deviations from the mean value in the first run.

The ratio between expected values (theoretical) to reported values, the ratio between measured to calculated conductivity and the ratio between equivalent concentrations of anions to equivalent concentrations cations, are presented in tables.

## Youden plot

The Youden plot is a graphical method to analyse inter-laboratory data where the samples are ordered in pairs with similar concentrations. One plot is made for each pair of samples and gives results for all participating laboratories. The plots visualize both systematic and random errors.

The plot is draw as a scatter plot where each point represents a pair of concentrations for one laboratory. The expected values for the two samples are drawn as solid blue lines. The arithmetic average of the measured values excluding outliers is drawn as dotted lines. The solid lines divide the plot in four quadrants and a 45 reference line going through the intercept of the solid lines may be added.

If errors are due to random factors, the points will be evenly distributed around the mean value and situated in all four quadrants.

If systematic errors dominate, the results will be close to 45 reference line, and be situated in the upper right quadrant (overestimation) or lower left quadrant (underestimation).

Drawing a line from a given point perpendicular on the 45 reference line gives two line segments, one from the point to the intercept on the reference line (a), and one continuing from the intercept to the point representing the expected values (b). The lengths of these line segments are measures of the random and systematic errors respectively.

Ellipses with radii corresponding to the data quality objectives (DQO, Table 1) are added in each plot. The data points are colour coded depending on the magnitude of errors as given in Table 1.

Figure 1:	Youden plot showing concentrations for a pair of samples (green), expected values (blue lines), average of measured values (dotted lines) and random and systematic errors (red lines)

In Figures 18–33 the reported data are presented in Youden plots.

Table 1:	Youden plot parameters.

| Radii | Components |
| --- | --- |
| 10% | SO2 in abs.sol, NO2 in abs.sol. |
| 20% | SO2, HNO3 and NH3 in impregnated filter |
| Radii = DQO | Components |
| 10% accuracy or better | SO4 2- -, NO3- |
| 15% accuracy or better | NH4+, Cl-, Ca2+, K+, Mg2+, Na+ , cond, H+ (from pH) |
| 0.1 units | pH |
| Criteria | Colour |
| Within 0.5*DQO | Blue |
| Within DQO | Green |
| Within 2*DQO | Orange |
| > 2*DQO | Red |

## Estimating random and systematic errors from laboratory comparisons

Table 12 presents relative random and relative systematic errors obtained by the different laboratories in the analysis of each parameter in the precipitation samples. The calculation method and assumptions used are given in Chapter 3.3.1 and Chapter 3.3.2.

### Estimating random errors

Systematic errors or bias in the laboratory analyses give a constant shift in the results from the expected ones at a particular concentration level. It is assumed that laboratories taking part in comparisons will obtain results near the expected ones when this bias is removed, and that the differences between expected and obtained results more often will be close to zero than not. A triangular distribution, based upon this assumption, can be used to quantify the random errors in the laboratory results (Eurachem/CITAC, 2000).

-a	0	a

x

The triangle distribution is symmetric with a baseline 2a. The height in the triangle will be 1/a when the triangle area equals 1. The standard uncertainty is given by

u(x) = a6	(1)

and more than 95 % of the data will be within ± 2∙u(x). The distance from -a to a (i.e. 2a) is called the range. When applied on the laboratory comparison results, the range equals the distance between the largest and smallest of the four differences between expected and found concentrations. As long as the bias can be assumed to be constant for the samples in the comparison of a specific component, it cannot have an effect on the distance corresponding to 2a. The bias may be dependent upon the concentrations, but can be considered approximate constant for the concentrations used here in the comparison of the main components in precipitation, since the differences between the concentrations are small.

L and T represent the laboratories’ and the expected concentrations respectively, and D is the difference. The difference for the lowest concentration is

D1 = L1 – T1	(2)

and the differences are D1, D2, D3, D4 in increasing order.

The range is D4 – D1 and the standard uncertainty for the differences u(D) becomes

u(D) = D4 − D12 ⋅ 6.    	(3)

The average expected concentration T for the four samples is given by

T = T1 + T2 + T3 + T44	(4)

The relative standard uncertainty, RSD, for 4 samples is given by uDT, or

RSD = 2 ⋅ D4 − D1 ⋅ 1006 ⋅ T1 + T2 + T3 + T4%,	(5)

and 95 per cent of the laboratory results in this comparison are expected to be within ±2 · RSD.

If the data quality objectives (DQO) likewise are looked upon as 95 percentiles, then 95 per cent of the laboratory analytical results should not be more than 10 or 15 per cent from the correct values (10 per cent for S and N containing components and 15 per cent for other components).

Correspondingly, the values 2·RSD should therefore be less than 10 or 15 per cent in order to comply with the DQO.

### Estimating systematic errors

An estimation of bias in single measurements requires a long data series, and four samples as we normally have in laboratory comparison, are merely able to give an indication of the bias or a very coarse estimate.

Coarse estimates have been performed here in the cases where the four samples had similar concentrations and where all four laboratory results were either higher or lower than the expected concentrations. The median of the differences Di, as defined above, was taken as a measure of the bias, B, in these cases.

B = median[Di]	(6)

A relative bias, RB, was also calculated based upon the average expected concentration T, as defined in (4).

RB = 4 ⋅ medianDi ⋅ 100T1 + T2 + T3 + T4%	(7)

# Results

## Sulphur dioxide in absorbing solution (A-samples)

Four samples and one blank solution were distributed to the laboratories that use the hydrogen peroxide absorption solution method. For those laboratories that reported a blank value this has been subtracted from the reported results.

7 laboratories have reported values for SO2 in absorbing solution. 39% of the reported values are within 10% of expected value. One laboratory reports 4 values below 20% of expected value.

The results are presented in Table 13 and Figure 2.

## Sulphur dioxide and nitric acid on impregnated filter (B-samples)

Five impregnated filter samples (including one blank) for determination of sulphur dioxide were analysed by 20 laboratories. The value reported for the blank filter was subtracted from the other values before the data were used.

The amount of sulphur on the distributed filters corresponds to air concentrations between 0.64-3.04 µg S m-3 when 25 m3 is sampled.

In addition to sulphate, nitrate was added to the same impregnated filters for determination of HNO3-N. The value reported for the blank filter was subtracted from the other values before using the data.

The amount of nitrogen on the distributed filters corresponds to air concentrations between 0.46 µg N m-3–1.05 µg N m-3 when 25 m3 sampling volume is used.

50% of the reported results for sulphur dioxide and 72% of the reported results for nitric acid are within 10% of expected value.

The results are presented in Tables 14 and 15 and Figures 3 and 4.

## Nitrogen dioxide in absorbing solution (C-samples)

The four samples distributed were made to represent both absorption solutions and extracts from iodide-impregnated glass filters. The samples contain known amounts of sodium nitrite diluted in water. In order to assure sample stability and to give the laboratories the opportunity to use the matrix they use in their daily routine, the distributed samples were to be diluted 1:10. The results should be reported as the diluted concentrations.

The 10 times diluted samples correspond to air concentrations between 1.85-5.95 µg NO2-N m-3, when 70 ml absorbing solution and 1.4 m3 are used. If 4 ml extraction solution and 0.7 m3 sampling volume are used, the samples correspond to air concentrations between 0.21-0.68 µg NO2-N m-3.

19 laboratories have reported results. 74% of the reported results are within 10% of expected value. Five laboratories report results that deviates more than 20% from expected value. Of these results all except for one, are above expected value.

The results are presented in Table 16 and Figure 5.

## Ammonia on impregnated filters (J-samples)

Six impregnated filters inclusive two unidentified blank filters were sent to 25 laboratories. 20 laboratories have reported their analytical results. The two blank values reported by each laboratory were averaged and subtracted from the other values reported before the data were used. The results are shown in Table 17 and Figure 6.

The amount of nitrogen on the filters correspond to air concentrations between 0.44-1.12 µg N m-3, if 25 m3 sampling volume is used.

Nine laboratories report all results within 10% of expected value. Three laboratories report values outside DQO.

## Precipitation (G-samples)

Four precipitation samples were distributed and 2474 single results from 68 laboratories were reported. 94 results were identified as outliers. This is < 4% of the data, which is a decrease compared to earlier rounds.

Results for sulphate are mostly in good agreement with expected value and only 4 laboratories report results outside DQO.  Standard deviation when outliers are excluded varies between 5-6%, which is very good.  The reported results for nitrate are overall in good agreement with expected value, though an increase in reported results outside DQO is observed when compared to last year’s results. Standard deviation for nitrite varies between 14-17% when outliers are excluded.

Results for magnesium and potassium shows an improvement compared to last year.  Fewer results outside DQO are reported.

Determination of calcium still seems to be difficult for many laboratories. 40% of the laboratories reports results outside DQO.

The results are presented in Tables18-28 and in Figures 7-17.

# Summary

A total of 68 laboratories participated in the twenty-sixth intercomparison. 37 of these laboratories are within the EMEP network.

As in earlier intercomparisons, outliers are defined as values that deviates more than two standard deviations from the mean value. Outliers occur for all samples and parameters. Out of a total of 2818 single results, 111 are defined as outliers. This is less than 4% of the reported data, which is very good.

# References

EMEP (1996) Manual for sampling and chemical analysis. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 1/95) (Newest revision in 2001).

Eurachem/CITAC (2000) Quantifying uncertainty in analytical measurements. 2nd ed. URL: http://www.measurementuncertainty.org/mu/QUAM2000-1.pdf.

Hanssen, J.E. (1988) The tenth intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 3/88).

Hanssen, J.E. (1990) The eleventh intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 1/90).

Hanssen, J.E. and Ladegård, N.E. (1984) The seventh intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 1/84).

Hanssen, J.E. and Ladegård, N.E. (1985) The eighth intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 4/85).

Hanssen, J.E. and Ladegård, N.E. (1987) The ninth intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 4/87).

Hanssen, J.E., Ladegård, N.E. and Thrane, K.E. (1983) The sixth intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 2/83).

Hanssen, J.E. and Skjelmoen, J.E. (1992) The twelfth intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 7/92).

Hanssen, J.E. and Skjelmoen, J.E. (1994) The thirteenth intercomparison of analytical methods within EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 1/94).

Hanssen, J.E. and Skjelmoen, J.E. (1995) The fourteenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 3/95).

Hanssen, J.E. and Skjelmoen, J.E. (1996) The fifteenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 2/96).

Hanssen, J.E. and Skjelmoen, J.E. (1997) The sixteenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 2/97).

Hanssen, J.E. and Skjelmoen, J.E. (2001) The seventeenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 10/2001).

Hjellbrekke, A.-G., Uggerud, H., Hanssen, J.E. and Schaug, J. (2005) The twenty-second intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 8/2005).

Thrane, K.E. (1978) Report on the first intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 2/78).

Thrane, K.E. (1980a) Report on the second and third intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 1/80).

Thrane, K.E. (1980b) Report on the fourth intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 5/80).

Thrane, K.E. (1981) The fifth intercomparison of analytical methods within the EMEP. Lillestrøm, Norwegian Institute for Air Research (EMEP/CCC-Report 2/81).

Uggerud, H.T., Hanssen, J.E. and Skjelmoen, J.E. (2001) The eighteenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 11/2001).

Uggerud, H.T., Hanssen, J.E., Schaug, J. and Skjelmoen, J.E. (2002) The nineteenth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 1/2002).

Uggerud, H.T., Hanssen, J.E., Schaug, J. and Skjelmoen, J.E. (2003) The twentieth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 8/2003).

Uggerud, H.T., Hanssen, J.E., Schaug, J. and Hjellbrekke, A.-G. (2004) The twenty-first intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 6/2004).

Uggerud, H. and Hjellbrekke, A.-G. (2007) The twenty-third and twenty-fourth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 5/2007).

Uggerud, H. and Hjellbrekke, A.-G. (2008) The twenty-fifth intercomparison of analytical methods within EMEP. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 7/2008).

- Tables

Table 2:	Samples distributed for the twenty-second interlaboratory test.

| A. | 5 synthetic samples for determination of SO2, consisting of 0.3% H2O2 absorbing solution and containing different concentrations of sulphuric acid. One of the samples was an unidentified blank. |
| --- | --- |
| B. | 5 KOH-impregnated Whatman 40 filters, comprising 1 blank and 4 filters to which different amounts of sulphuric acid and nitrate salt have been added. |
| C. | 4 synthetic samples for determination of NO2 consisting of sodium nitrite diluted in water. |
| J. | 6 Whatman 40 filters impregnated with 3% oxalic acid, comprising 2 blank and 4 filters to which different amounts of ammonium salt solution have been added. |
| G. | 4 synthetic precipitation samples, containing SO4²-, NO3-, NH4+, H+, Na+, Mg2+ and Cl-, and Ca2+ and K+. |

Table 3a:	EMEP laboratories participating in the twenty-second laboratory intercomparison. The numbers in front of the names are used in tables and figures.

| Armenia | (45) | Dept of Environm. Health and Research |
| --- | --- | --- |
| Austria Belgia | ( 1) (2) | Umweltbundesamt Zweigstelle Sud, Klagenfurt Flemish Environment Agency, Antwerpen |
| Canada | (26) | Meteorological Service of Canada, Toronto |
| Croatia | (35) | Meteorological and Hydrological Service of Croatia |
| Czech Republic | ( 3) | Czech Hydrometeorological Institute, Praha |
| Denmark | ( 4) | National Environmental Research Institute. Air Pollution Laboratory |
| Estonia | (38) | Estonian Environmental Research Centre, Tallinn |
| Finland | ( 5) | Finnish Meteorological Institute. Air Quality Department |
| France | (41) | Micro Pollutants Technology |
| Georgia | (43) | Centre for Monitoring and Prognostication, Tbilisi |
| Germany | ( 7) | IfE Leipzig GmbH, Umweltlabor |
| Germany | ( 8) | Umweltbundesamt, Messtelle Schauinsland |
| Hungary | (10) | Institute for Atmospheric Physics |
| Ireland | (12) | Met. Eirann, Dublin |
| Italy | (13) | C.N.R. Istituto Inquinamento Atmosferico |
| Italy | (30) | Joint Research Centre, Ispra |
| Latvia | (33) | Air Pollution Observation Laboratory |
| Lithuania | (32) | Environmental Physics and Chemistry Laboratory |
| Netherlands | (14) | National Institute of Public Health and Environmental Protection (RIVM) |
| Norway | (15) | Norwegian Institute for Air Research (NILU) |
| Macedonia | (40) | Hydrometeorological Institute, Skopje |
| Moldova | (42) | State Hydrometeorological Service, Chisinau |
| Poland | (16) | Institute of Meteorology and Water Management, Warsaw |
| Poland | (39) | Environmental Monitoring Laboratory, Institute of Environmental Protection |
| Poland | (46) | Diabla Gora station |
| Portugal | (17) | Laboratorio Santo Andre |
| Russian Federation | (22) | Institute of Global Climate and Ecology |
| Serbia and Montenegro | (24) | Rep. Hydrometeorological Institute of Serbia |
| Slovakia | (31) | Slovak Hydrometeorological Institute |
| Slovenia | (36) | Hydrometeorological Institute of Slovenia |
| Spain | (19) | Centro Nacional de Sanidad Ambiental |
| Sweden | (20) | Swedish Environmental Research Institute (IVL), Gothenburg |
| Switzerland | (21) | Swiss Federal Laboratories for Materials Testing (EMPA) |
| Turkey | (34) | Refik Saydam Hygiene Center, Ankara |
| United Kingdom | (23) | AEA Technology, National Environmental Technology Centre |
| United States of America | (27) | Illinois State Water Survey |

- Table 3b:	Participating laboratories outside the EMEP network.

| Germany | (104) | Hessige Landwirtschafliche |
| --- | --- | --- |
| Finland | (107) | The Finnish Forest Institute |
| Germany | (108) | Institut f. Bondenkunde und Standortlehre, Dredsten |
| Germany | (109) | Bügen-Institute, dep. of Soil Science of Temporal and Boreal Ecosystems |
| Germany | (110) | Thüringer Landesanstalt für Landwirtschaft (TTL), Jena |
| Germany | (112) | Niedersächsische Forstliche Versuchsanstalt (NVF) |
| Germany | (113) | Landesforstanstalt Eberswalde, abt. Waldökologie |
| Italy | (114) | C.N.R. Istituto Italiano di Idrobiologia |
| Germany | (115) | Bayerische Landesanstalt f. Wald- und Forstwirtshaft |
| Switzerland | (116) | Institute for Applied Plant Biology |
| Germany | (118) | Forstliche Versuchs-und Forschungsansta |
| Germany | (120) | Landwirtschaftliche Untersuchungs- und Forschungsanstalt (LUFA) |
| Germany | (121) | Landeslabor Schleswig-Holstein |
| Belgium | (124) | Laboratorium voor Bondemkunde, Gent |
| Germany | (125) | Bayerisches Landesamt für Umweltschutz, Augsburg |
| Italy | (126) | APPA Laboratorio Biologico Provinciale |
| Belarus | (133) | Institute for Problems of Natural Resources Use and Ecology, Minsk |
| Italy | (140) | C.N.R. Istituto di Ricerca sulle Acque |
| Finland | (145) | Tartu Environmental Research, Tartu |
| Spain | (150) | Fundación Centro de Estudios ambentials del meditirrain |
| Belgium | (151) | Laboratoire de l'Unitè des Eaux et Forêst (EFOR), Louvain-la-Neuve |
| Norway | (152) | Norwegian Forest Research Institute, Ås |
| Slovenia | (153) | Slovenian Forestry Institute, Ljubljana |
| United Kingdom | (155) | Environmental Research Branch, Forest Research |
| Hungary | (157) | Ecological Laboratory of Forest research Institute |
| Japan | (158) | Acid Deposition and Oxidant Research Center (ADOCRC), Niigata |
| Thailand | (163) | Environmental Researching and Training Center (ERTC) |
| Thailand | (164) | Pollution Control Department (PCD) |
| Viet Nam | (165) | Institute of Meteorology and Hydrology |
| Poland | (166) | Forest Research Institute, Laboratory of Forest Habitat Chemistry |
| United Kingdom | (167) | CEH Edinburgh |

Table 4:	Analytical methods used at the participating laboratories for the determination of sulphur dioxide in absorbing solution (A).

| Method | Laboratory |
| --- | --- |
| - Ion chromatography - Spectrophotometry | 15, 17, 19, 23, 34, 36 173 |

Table 5:	Analytical methods used at the participating laboratories for the determination of sulphur dioxide on impregnated filters (B).

| Method | Laboratory |
| --- | --- |
| - Ion chromatography | 3, 4, 5, 15, 20, 31, 32, 33, 34, 36, 38, 41, 158, 163, 164, 165 |
| - Capillary Electrophoresis - Spectrophotometry | 39 16 |

Table 6:	Analytical methods used at the participating laboratories for determination of nitric acid on impregnated filters (B).

| Method | Laboratory |
| --- | --- |
| 1. Ion chromatography | 3, 4, 5, 15, 20, 31, 32, 33, 34, 36, 38, 41, 158, 163, 164, 165, |
| 2. Capillary Electrophoresis 3. Spectrophotometry | 39 16 |

Table 7:	Analytical method for determination of ammonia on impregnated filters (J).

| Method | Laboratory |
| --- | --- |
| 1. Spectrophotometry | 3, 4, 8, 10, 15, 16, 20, 32, 33, 34, 39, 45 |
| 3 Ion chromatography | 5, 13, 31, 36, 41, 158, 165 |

Table 8:	Analytical method used for NO2 in absorbing solution (C).

| Method | Laboratory |
| --- | --- |
| 1. Spectrophotometry | 3, 4, 8, 10,12, 15, 16, 20, 23, 31, 32, 33, 34, 35, 36, 38, 39, 45, 173 |
| 2. Ion chromatography | 19 |

Table 9:	Analytical methods used for the determination of chemical constituents in precipitation samples.

| Lab no | Network | SO42- | NH4+ | NO3- | Na+ | Mg2+ | Cl- | Ca2+ | K+ |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 2 | Select network | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 3 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 4 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AES | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AES |
| 5 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 7 | Select network | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 8 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 10 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 12 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 13 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 14 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Na methods, ICP-MS | Mg methods, ICP-MS | Ion chromatography | Ca methods, ICP-MS | K methods, ICP-MS |
| 15 | EMEP | Ion chromatography | Ion chromatography, Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 16 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 17 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 19 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 21 | 0.003 | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 22 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 23 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 24 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 26 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 27 | Other | Ion chromatography | Spectrophotometry | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 30 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 31 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 32 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Mg methods | Ion chromatography | Flame-AAS | Flame-AAS |
| 33 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AES | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AES |
| 34 | Select network | Ion chromatography | Spectrophotometry | Ion chromatography | Na methods | Flame-AAS | Ion chromatography | Flame-AAS | K methods |
| 35 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 36 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 38 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 39 | EMEP | SO4 methods, CE | Spectrophotometry | NO3 methods, CE | ICP-AES, PN-ICP-AES(A) | ICP-AES, PN-ICP-AES(A) | Cl methods, CE | ICP-AES, PN-ICP-AES(A) | ICP-AES, PN-ICP-AES(A) |
| 40 | EMEP | SO4 methods | Spectrophotometry | Spectrophotometry | Flame-AAS | Flame-AAS | Spectrophotometry | Flame-AAS | Flame-AAS |
| 41 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 42 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 43 | Select network | Ion chromatography | NH4 methods, Photocolorimeter | Ion chromatography | Na methods | Mg methods | Ion chromatography | Ca methods | K methods |
| 45 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Na methods, ICP-MS | Mg methods, ICP-MS | Ion chromatography | Ca methods, ICP-MS | K methods, ICP-MS |
| 104 | ICP-Forest | Ion chromatography | NH4 methods, Photocolorimeter, CFA | NO3 methods, CFA | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 108 | 0.04 | Ion chromatography | Spectrophotometry, FIA | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 109 | EMEP | ICP-AES | NH4 methods, Continuous flow | NO3 methods, Continuous flow | ICP-AES | ICP-AES | Cl methods, Continuos flow. electrochemical potential of AgCl | ICP-AES | ICP-AES |

Table 9, cont.

| Lab no | Network | SO42- | NH4+ | NO3- | Na+ | Mg2+ | Cl- | Ca2+ | K+ |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 110 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 112 | ICP-Forest | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | ICP-AES | Ion chromatography | ICP-AES | Ion chromatography |
| 113 | Other | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | ICP-AES | Ion chromatography | ICP-AES | Ion chromatography |
| 114 | ICP-Forest | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 115 | ICP-Forest | Ion chromatography | Ion chromatography | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 118 | EMEP | Ion chromatography | NH4 methods, Continuous flow indophenol blue | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 120 | ICP-Forest | SO4 methods | NH4 methods | NO3 methods | Na methods | Mg methods | Cl methods | Ca methods | K methods |
| 121 | EMEP | Spectrophotometry | Spectrophotometry | Spectrophotometry | ICP-AES | ICP-AES | Spectrophotometry | ICP-AES | ICP-AES |
| 125 |  | Ion chromatography | Ion chromatography | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 126 | ICP-Forest | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 133 | EMEP | SO4 methods, Turbidimetry | NH4 methods, Colimetry | NO3 methods, Colorimetry | Na methods, Flame Photometry | Mg methods, Atomic Absorption | Cl methods, Titration | Ca methods, Atomic absorption | K methods, Flame Photometry |
| 140 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 145 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 150 | ICP-Forest | SO4 methods | NH4 methods | NO3 methods | Na methods | Mg methods | Cl methods | Ca methods | K methods |
| 151 |  | Ion chromatography | Spectrophotometry | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 152 | ICP-Forest | Ion chromatography | Spectrophotometry | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 153 | ICP-Forest | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 155 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 157 | EMEP | Ion chromatography | Spectrophotometry | Ion chromatography | Flame-AAS | Flame-AAS | Ion chromatography | Flame-AAS | Flame-AAS |
| 158 | EANET | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 163 | EMEP | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |
| 164 | EMEP | SO4 methods | NH4 methods | NO3 methods | Na methods | Mg methods | Cl methods | Ca methods | K methods |
| 165 | EANET | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Flame-AAS | Ion chromatography | Flame-AAS | Ion chromatography |
| 166 | ICP-Forest | Ion chromatography | Ion chromatography | Ion chromatography | ICP-AES | ICP-AES | Ion chromatography | ICP-AES | ICP-AES |
| 167 | EMEP | Ion chromatography | NH4 methods, Ammonia Flow Injection analyser | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography | Ion chromatography |

Table 10:	Reported results for precipitation samples expressed as % deviation from expected value.

| Precipitation |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Lab no | SO42- |  |  |  | NH4+ |  |  |  | NO3- |  |  |  | Na+ |  |  |  | Mg2+ |  |  |  | Cl- |  |  |  | Ca2+ |  |  |  | K+ |  |  |  | pH |  |  |  | Cond |  |  |  | H+ |  |  |  |
|  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | pH-units from expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  |
|  | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 |
| 1 | 1 | 2 | 1 | 1 | 6 | 8 | 3 | 2 | 0 | 0 | 0 | 1 | -2 | -3 | -3 | -2 | -11 | -15 | -11 | -13 | -1 | -2 | -1 | -2 | -28 | -22 | -19 | -17 | -7 | -7 | -8 | -5 | 0.03 | 0.02 | 0.05 | 0.02 | -7 | -3 | 0 | 0 |  |  |  |  |
| 2 | 6 | 8 | 9 | 10 | 7 | 9 | 8 | 6 | 7 | 10 | 9 | 10 | -4 | -6 | -4 | 1 | 10 | 29 | 25 | 16 | 31 | 37 | 24 | 19 | -3 | 11 | 3 | 0 | 14 | 17 | 22 | 11 | -0.2 | -0.19 | -0.28 | -0.29 | -3 | 0 | -2 | 0 |  |  |  |  |
| 3 | 1 | 1 | 1 | 0 | 30 | 31 | 22 | 17 | 0 | 1 | -1 | 0 | -2 | -4 | -5 | -6 | 1 | 3 | 1 | -2 | -1 | -3 | -2 | -3 | -2 | -6 | -10 | -7 | -1 | 3 | -1 | 1 | -0.013 | -0.025 | -0.029 | -0.013 | 0 | 4 | 0 | 3 |  |  |  |  |
| 4 | 0 | 1 | 0 | -1 | -4 | -1 | 0 | 0 | -4 | -4 | -3 | -1 | 8 | 5 | 4 | 5 | -12 | -22 | -18 | -18 |  |  | 7 | -9 | -2 | -1 | -1 | -2 | -12 | -9 | -8 | -8 | 0.04 | 0.04 | 0.04 | 0.03 | -3 | 5 | 2 | 5 |  |  |  |  |
| 5 | 0 | 0 | -1 | 0 | -2 | -1 | 0 | 0 | 2 | 2 | 1 | 2 | 5 | 3 | 3 | 3 | 2 | 0 | 3 | -1 | 3 | 2 | 1 | 1 | 2 | 9 | 3 | 2 | 6 | 6 | 4 | 2 | 0.01 | 0 | -0.01 | -0.02 | 5 | 6 | 5 | 8 |  |  |  |  |
| 7 | 0 | -1 | -3 | -3 | -2 | -1 | 2 | 3 | -2 | -2 | -2 | -1 | 6 | 6 | 6 | 6 | 1 | -6 | 3 | -4 | 14 | 10 | 1 | 0 | 31 | 10 | 11 | 6 | 2 | 1 | 2 | 2 | 0.03 | 0.04 | 0.06 | 0.04 | -16 | -13 | -11 | -7 |  |  |  |  |
| 8 | 2 | 4 | 5 | 4 | -5 | -2 | 0 | -2 | 4 | 6 | 6 | 6 | 0 | -1 | -1 | 0 | 4 | 4 | 5 | 4 | 7 | 3 | -1 | -2 | 20 | 10 | 9 | 9 | -7 | -6 | -3 | -1 | 0.04 | 0.04 | 0.04 | 0.04 | -5 | -2 | -4 | -1 |  |  |  |  |
| 10 | -5 | -4 | -3 | -4 | -1 | -3 | 1 | 2 | -5 | -3 | -4 | -2 | -11 | -12 | -11 | -10 | 5 | 10 | 7 | 4 | -11 | -11 | -11 | -12 | 45 | 20 | 16 | 15 | 15 | 15 | 14 | 13 | 0.12 | 0.08 | 0.06 | 0.07 | 0 | 1 | 0 | 2 |  |  |  |  |
| 12 | -7 | -6 | -5 | -6 | 2 | 6 | 5 | 3 | -6 | -5 | -3 | -3 | -5 | -5 | -4 | -3 | 2 | -11 | -5 | -3 | -9 | -8 | -7 | -6 | -22 | -16 | -17 | -13 | -7 | -4 | -8 | -8 | 0.02 | 0.02 | 0 | 0.01 | 1 | 4 | 0 | 2 |  |  |  |  |
| 13 | 6 | 10 | 16 | 9 | -3 | 0 | 1 | -2 | 7 | 10 | 10 | 14 | 3 | 2 | 2 | 0 | -2 | -1 | -2 | -8 | 13 | 9 | 11 | 16 | -4 | -4 | -3 | 0 | -5 | -3 | -2 | -8 | 0.07 | 0.01 | 0.03 | 0 | -3 | 0 | 5 | 30 |  |  |  |  |
| 14 | -1 | -2 | -3 | -2 | 3 | 4 | 0 | -2 | 1 | 2 | 0 | -1 | 0 | 3 | -1 | 1 | -10 | -6 | -6 | -11 | -27 | -26 | -24 | -11 | 2 | -4 | 4 | -4 | -8 | -13 | -8 | -13 | 0.1 | 0.11 | 0.13 | 0.14 | 4 | 7 | 4 | 7 |  |  |  |  |
| 15 | 0 | 5 | 4 | 3 | 5 | 3 | -3 | -4 | -2 | -3 | -5 | -3 | -3 | -3 | -5 | -3 | 18 | 31 | 25 | 10 | 2 | -5 | -8 | -9 | 0 | -15 | -5 | -6 | -1 | -4 | 1 | -3 | 0.06 | 0.04 | 0.05 | 0.03 | -1 | 3 | 0 | 3 |  |  |  |  |
| 16 | -3 | -2 | -2 | -3 | 9 | 3 | -2 | 2 | -5 | -6 | -4 | -3 | 3 | -1 | -4 | -3 | -4 | 4 | 0 | 2 | -4 | -7 | -5 | -3 | 0 | 2 | 1 | 4 | -7 | -4 | -3 | -1 | 0.03 | 0.02 | 0.01 | 0.01 | -4 | -3 | -2 | 0 |  |  |  |  |
| 17 | -21 | -19 | -18 | -19 | 16 | 19 | 12 | 12 | -15 | -15 | -14 | -10 |  | -81 | -63 | -58 | -17 | -39 | -31 | -31 | 203 | 97 | 29 | 101 | -33 | -13 | -15 | -13 | -20 | -32 | -39 |  | -0.08 | -0.06 | -0.03 | -0.1 | -6 | 1 | -2 | 2 |  |  |  |  |
| 19 | -11 | -8 | -7 | -7 | -1 | -1 | -10 | -9 | -11 | -9 | -8 | -7 | 6 | 6 | 1 | 2 | 1 | -3 | 2 | 0 |  |  | 7 |  | 4 | 8 | 7 | 2 | 3 | 4 | 11 | 11 | 0.14 | 0.13 | 0.15 | 0.15 | 0 | 3 | -1 | 3 |  |  |  |  |
| 20 | -3 | -3 | -1 | -2 | 9 | 13 | 5 | 6 | 1 | -2 | 0 | -1 | 19 | 9 | 9 | 9 | 31 | 39 | 25 | 14 | -3 | -2 | 0 | 0 | 275 | 33 | 15 | 13 | 8 | 2 | 5 | 9 | 0.04 | 0.02 | 0.01 | 0 | 1 | -43 | 2 | 4 |  |  |  |  |
| 21 | 1 | 3 | 1 | 2 | -2 | -1 | 1 | 1 | -1 | 0 | 0 | 1 | 2 | 5 | -1 | -1 | -2 | -7 | -3 | -5 | -8 | -7 | -5 | -3 | 1 | -1 | -4 | 0 | -1 | -2 | -4 | -5 | 0.04 | 0.03 | 0.03 | 0.02 | 1 | 4 | 0 | 4 |  |  |  |  |
| 22 | 1 | 3 | -1 | 2 | 1 | -2 | -5 | -6 | 7 | 19 | 16 | 17 | -3 | -12 | -11 | -15 | 2 | 4 | 5 | -2 | 27 | 13 | 3 | -2 | 78 | 14 | -7 | -32 | 5 | -13 | -22 | -26 | 0.03 | 0.02 | 0.06 | 0.08 | -18 | -21 | -18 | -22 |  |  |  |  |
| 23 | -5 | -4 | -4 | -3 | 10 | 10 | 9 | 9 | 0 | 1 | 0 | 2 | -1 | -2 | 0 | 2 | 7 | 25 | 16 | 7 | -5 | -4 | -2 | -1 | 25 | 18 | 10 | 5 | -4 | -2 | 1 | 8 | 0.12 | 0.08 | 0.06 | 0.04 | -9 | -1 | -5 | -4 |  |  |  |  |
| 24 | -4 | -8 | -7 | -7 | 3 | -7 | -2 | 5 | -7 | -11 | -10 | -8 | 3 | 7 | -6 | -5 | 9 | 17 | 7 | 6 | -75 | -75 | -66 | -60 | 0 | 12 | 4 | 3 | -4 | -5 | -7 | -14 | -2 | -2 | -2 | -2 | 6 | 6 | 0 | 6 |  |  |  |  |
| 26 | -4 | -3 | -2 | -2 | 2 | 1 | 0 | 2 | 4 | 4 | 0 | 0 | 1 | 0 | -1 | 0 | -2 | -1 | 0 | -4 | 2 | -4 | 1 | 1 | 4 | 6 | 2 | 3 | -2 | -2 | -3 | -1 | 0.07 | 0.06 | 0.09 | 0.11 |  |  |  |  |  |  |  |  |
| 27 | 3 | 5 | 5 | 5 | -1 | 0 | 0 | 1 | 3 | 3 | 4 | 5 | 2 | 0 | 1 | 1 | 2 | 1 | 3 | 0 | 1 | 0 | 1 | 2 | 3 | 3 | 2 | 1 | -2 | -2 | -2 | -3 | 0.06 | 0.06 | 0.03 | 0.05 | -4 | -1 | -2 | 0 |  |  |  |  |
| 30 | -2 | -1 | -3 | -1 | -7 | -1 | -3 | -1 | -12 | -11 | -10 | -8 | 5 | 3 | 4 | 4 | -2 | -6 | 0 | -3 | 10 | -2 | -4 | 1 | 2 | 1 | 4 | 3 | -8 | -7 | -8 | -8 | 0.12 | 0.13 | 0.09 | 0.03 | -6 | 2 | 0 | 0 |  |  |  |  |
| 31 | -2 | -2 | -3 | -2 | 26 | 26 | 15 | 14 | -3 | -2 | -2 | -1 | -6 | -7 | -7 | -7 | 1 | 3 | -1 | -4 | -1 | -2 | -3 | -4 | -1 | -10 | -14 | -13 | -6 | 3 | -8 | -6 | 0.07 | 0.08 | 0.04 | 0.04 | -1 | 2 | -2 | 3 |  |  |  |  |
| 32 | -2 | -1 | -1 | 0 | 3 | -1 | -3 | -2 | -1 | 1 | 0 | 1 | -2 | -1 | -2 | 2 |  |  |  |  | 6 | 3 | -2 | 2 | -41 | -28 | -6 | -13 | -2 | 2 | 1 | -2 | 0.03 | 0.02 | 0.02 | 0.02 | -4 | -1 | -4 | 0 |  |  |  |  |
| 33 | -2 | -3 | -4 | -3 | 4 | 5 | 0 | 3 | 1 | 2 | 0 | 3 | -5 | -6 | -1 | -7 | -6 | -1 | 0 | -2 | 4 | 3 | -3 | -3 | -3 | -1 | 2 | -4 | -8 | -5 | -5 | -4 | 0.06 | 0.06 | 0.07 | 0.06 | -2 | 1 | -1 | 1 |  |  |  |  |
| 34 | -9 | -9 | -10 | -10 | 1 | -3 | 9 | 15 | -13 | -14 | -14 | -13 |  |  |  |  | -10 | -8 | -4 | -3 | -82 | -61 | -68 | -51 | 14 | 13 | 18 | 68 |  |  |  |  | 0.143 | 0.091 | 0.091 | 0.069 | -10 | -7 | -4 | -5 |  |  |  |  |
| 35 | -1 | -1 | 0 | 0 | 12 | 14 | 14 | 14 | -3 | -3 | -1 | 1 | -14 | -16 | -10 | -3 | 14 | 18 | 19 | 18 | -9 | -5 | -8 | -8 | 11 | 33 | 42 | 34 | -36 | -39 | -47 |  | 0.07 | 0.05 | 0.06 | 0.08 | -4 | -1 | -6 | -6 |  |  |  |  |
| 36 | -3 | -2 | -3 | -3 | -11 | -12 | -7 | -6 | -5 | -5 | -6 | -5 | 0 | -2 | -2 | -2 | -6 | -6 | -3 | -5 | 23 | 13 | 9 | 7 | -2 | -1 | -1 | -2 | 0 | -6 | -11 | -11 | 0.04 | 0.04 | 0.02 | 0.03 | -11 | -10 | -10 | -7 |  |  |  |  |
| 38 | -2 | 7 | 7 | 7 | 1 | 2 | 1 | 0 | 4 | 2 | 0 | 3 | -34 | -34 | -34 | -31 | -15 | -14 | -13 | -15 | 17 | -2 | 0 | 4 | 11 | 8 | 5 | 10 | -20 | -23 | -22 | -19 | -0.12 | -0.12 | -0.11 | -0.16 | 1 | 3 | -4 | 3 |  |  |  |  |
| 39 | 13 | 13 | 7 | 8 | 10 | 10 | 10 | 9 | 13 | 15 | 7 | 10 | 9 | 7 | 8 | 9 | -1 | 0 | 1 | -2 | 16 | 16 | 8 | 12 | 2 | 2 | 1 | 2 | 0 | 0 | 0 | 2 | 0.09 | 0.07 | 0.07 | 0.09 | 1 | -1 | -1 | 6 |  |  |  |  |
| 40 |  |  |  |  | 7 | 9 | 6 | -1 | 23 | 18 | 26 | 31 | 27 | 30 | 15 | 52 | 10 | 269 | 700 | 231 | 1869 | 1643 | 980 | 925 | 45 | -57 | -45 | -54 | -3 | 7 | -14 | 60 | -0.16 | -0.22 | -0.22 | -0.26 | -11 | -11 | -10 | -4 |  |  |  |  |
| 41 | 3 | 1 | -1 | 0 | 25 | 27 | 23 | 20 | -5 | -5 | -2 | 0 | 6 | 15 | 4 | 7 | -1 | -19 | -11 | -10 | -18 | -18 | -20 | -22 | 17 | 22 | 11 | 6 | 3 | 8 | 9 | 11 |  |  |  |  |  |  |  |  |  |  |  |  |
| 42 | 34 | 36 | 33 | 38 | 29 | 36 | 33 | 37 | 47 | 50 | 48 | 51 | -17 | -18 | -20 | -13 | 5 | 0 | -2 | 9 | 55 | 62 | 17 | 43 | 8 | 0 | -4 | 11 | -5 | -11 | -13 | 3 |  |  |  |  |  |  |  |  |  |  |  |  |
| 43 | -7 | -7 | -11 | -6 | -20 | -14 | -30 | -30 | -83 | -87 | -84 | -83 |  |  |  |  |  |  |  |  | -11 | 169 | 38 | 106 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 45 | -3 | -2 | -4 | 0 | -2 | 8 | -11 | 0 | -1 | 2 | 0 | -7 | 3 | 6 | 0 | 1 | -2 | 2 | 0 | -1 | -9 | -6 | -4 | -10 | 1 | -2 | 0 | 3 | -6 | 1 | -5 | -3 | 0.005 | 0.01 | 0.005 | 0.01 | -7 | -6 | 10 | -3 |  |  |  |  |
| 46 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 0.07 | 0.06 | 0.04 | 0.05 | 0 | 0 | -1 | 4 |  |  |  |  |
| 104 | 19 | 17 | 13 | 12 | -32 | -25 | -14 | -7 | -2 | 2 | 0 | -1 | -8 | -15 | -14 | -13 | -2 | -24 | -1 | -6 | 17 | 26 | 21 | 17 | -5 | 2 | -7 | -8 | -9 | -10 | -1 | -13 | 0.61 | 0.13 | 0.02 | -0.01 | -37 | -11 | -5 | 0 |  |  |  |  |
| 107 | -2 | -4 | -6 | -5 | -3 | -1 | 0 | 1 | -9 | -7 | -7 | -6 | -3 | -4 | -4 | -4 | 7 | 24 | 27 | 18 | -15 | -16 | -17 | -18 | 7 | 72 | 69 | 57 | -7 | -5 | -8 | -9 | 0.04 | 0.04 | 0.02 | 0.01 | -5 | -1 | -5 | 1 |  |  |  |  |
| 108 | -4 | -3 | -1 | -2 | -12 | -6 | -6 | -7 | -2 | -2 | 0 | 3 | 7 | 6 | -1 | -3 | -14 | -31 | -32 | -21 | 114 | 84 | 93 | 75 |  |  | 30 | 22 | 3 | 2 | 5 | 9 | 0.11 | 0.11 | 0.09 | 0.11 | -11 | -10 | -11 | -5 |  |  |  |  |
| 109 | -14 | -17 | -16 | -16 | 36 | 13 | 1 | 12 | -6 | -2 | -3 | -1 | -30 | -38 | -35 | -3 | -3 | -3 | 2 | -4 |  |  |  | -16 | 41 | 14 | 20 | 4 |  |  |  |  | 0.17 | 0.14 | 0.11 | 0.32 | -3 | -3 | -5 | -2 |  |  |  |  |
| 110 | 3 | 1 | 3 | 12 | -18 | -25 | -14 | -13 | -1 | -3 | -4 | 3 | 10 | 3 | 3 | 5 | 1 | -28 | -19 | -17 |  |  |  |  | 46 | -8 | -9 | -6 |  |  |  |  | 0.18 | 0.14 | 0.24 | 0.14 | -9 | -6 | -7 | -3 |  |  |  |  |
| 112 | 0 | 2 | 1 | 1 | -6 | -5 | -1 | 0 | 10 | 9 | 5 | 4 | 6 | 4 | 3 | 3 | -5 | -10 | -6 | -8 | 1 | -1 | 19 | 17 | 2 | 1 | -1 | 1 | -1 | -4 | -3 | -6 | 0.12 | 0.11 | 0.12 | 0.15 | -10 | -7 | -9 | -7 |  |  |  |  |
| 113 | -2 | -1 | -5 | -3 | 13 | -1 | 4 | 31 | 6 | 2 | 4 | 1 | 0 | -5 | 0 | 3 | -6 | -8 | -9 | -12 | 12 | -21 | 14 | -6 | 17 | 17 | 4 | 24 | 0 | -3 | 27 | 60 | 0.08 | 0.07 | 0.1 | 0.09 | -13 | -7 | -14 | -9 |  |  |  |  |
| 114 | 2 | -1 | -4 | -3 | -4 | -1 | 1 | -3 | 0 | -1 | 0 | 1 | -1 | -1 | -1 | -1 | 14 | 11 | 25 | 14 | -9 | 9 | -7 | -6 | -22 | -22 | -22 | -26 | -22 | -9 | -8 | -8 | 0.04 | 0.01 | 0.01 | 0.02 | -6 | 0 | -1 | 0 |  |  |  |  |
| 115 | 2 | 5 | 7 | 4 | 3 | 8 | -4 | -1 | 7 | 7 | 0 | 1 | -6 | -11 | -8 | -9 | -13 | -14 | -13 | -15 | 42 | 20 | 4 | 4 | -6 | -5 | -8 | -8 | -22 | -34 | -29 | -29 | 0.13 | 0.14 | 0.12 | 0.12 | -4 | -3 | -5 | -3 |  |  |  |  |

Table 10, cont.

| Precipitation |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Lab no | SO42- |  |  |  | NH4+ |  |  |  | NO3- |  |  |  | Na+ |  |  |  | Mg2+ |  |  |  | Cl- |  |  |  | Ca2+ |  |  |  | K+ |  |  |  | pH |  |  |  | Cond |  |  |  | H+ |  |  |  |
|  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | pH-units from expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  |
|  | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 | G1 | G2 | G3 | G4 |
| 116 | 1 | -1 | -2 | -2 | 5 | 6 | 7 | 7 | 2 | -2 | -3 | -1 | 5 | 2 | -1 | -1 | -2 | -6 | -3 | -4 | -11 | -9 | -13 | -12 | -2 | -3 | -5 | -9 | 2 | 3 | 3 | 3 | 0.08 | 0.01 | -0.07 | -0.1 | -11 | 12 | -4 | 11 |  |  |  |  |
| 118 | -9 | -5 | 3 | 0 | -12 | -11 | -7 | -6 | 1 | 2 | 0 | 1 | -14 | -20 | -17 | -15 | -5 | -8 | -5 | -8 | -29 | -20 | -10 | -9 | 2 | -4 | -6 | -9 | -17 | -20 | -15 | -8 | 0 | -0.01 | -0.01 | -0.01 | -9 | -4 | -5 | -4 |  |  |  |  |
| 120 | -5 | -4 | -3 | -3 | -46 | -56 | -36 | -19 | -2 | -5 | -3 | -1 | 3 | 2 | 1 | 2 | -3 | -1 | -3 | -5 | 49 | 9 | -14 | -22 | 0 | 2 | -4 | 4 | -12 | 7 | -2 | -16 | 0.01 | 0 | 0 | 0 | 0 | 4 | 0 | 4 |  |  |  |  |
| 121 | -7 | -6 | -7 | -6 | 29 | 29 | 27 | 31 | 336 | 340 | 341 | 342 | -91 | -86 | -57 | -52 | -1 | -4 | -3 | -4 | 6 | 1 | -10 | -11 | 13 | 4 | 3 | 3 | 54 | 27 |  |  | 0.05 | 0.03 | 0.04 | 0.03 | -2 | 4 | 4 | 6 |  |  |  |  |
| 124 | 1 | -2 | -4 | -6 | 16 | 19 | 20 | 25 | 4 | -5 | 6 | 1 | 11 | 6 | 4 | 6 | -9 | 11 | 2 | 5 | -3 | -2 | 3 | 0 | 64 | 33 | 20 | 22 | -2 | -4 | -8 | 9 | 0.06 | 0.08 | 0.08 | 0.04 | -3 | -6 | -7 | 0 |  |  |  |  |
| 125 | -1 | 0 | -3 | -1 | 3 | 5 | 10 | 15 | 0 | -1 | -2 | -1 | 1 | 0 | -2 | -1 | 3 | 13 | 17 | 8 | -1 | 2 | 4 | -4 | 0 | -6 | -5 | -7 | 0 | 7 | 1 | 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 126 | -13 | -11 | -11 | -7 | -48 | -14 | -37 | -26 | -6 | -5 | -7 | -1 | 7 | 6 | 1 | 4 | -26 | -3 | -43 | -4 | -22 | -20 | -17 | -13 | 88 | 27 | 15 | 13 | -7 | -9 | -8 | -8 | 0.07 | 0.06 | 0.07 | 0.06 | 6 | 9 | 8 | 13 |  |  |  |  |
| 133 | -48 | 13 | -25 | -23 | 48 | 41 | 19 | -8 | -19 | -15 | -8 | -21 | -59 | -45 | -22 | -31 | 1 | -4 | -5 | -4 | 80 | 80 | 27 | 31 | 12 | -75 | -65 | -64 | -2 | 2 | 5 | 1 | 0.19 | 0.15 | 0.19 | 1.18 | -20 | -13 | -17 | -36 |  |  |  |  |
| 140 | -4 | -2 | -1 | -2 | 3 | 4 | 5 | 6 | -5 | -4 | 0 | 0 | 3 | 2 | -1 | -1 | -3 | -3 | 2 | -4 | -3 | -8 | -7 | -6 | 41 | 14 | 4 | 0 | 32 | 23 | 18 | 26 | 0.1 | 0.07 | 0.09 | 0.1 | -15 | -10 | -14 | -13 |  |  |  |  |
| 145 | -32 | -16 | -11 | -9 | 12 | 11 | 28 | -7 | -49 | -27 | -24 | -24 | 9 | 5 | -6 | 1 | 74 | 71 | 14 | -7 | -1 | 36 | 0 | 7 |  |  |  |  | -7 | -29 | -20 |  |  |  |  |  | -14 | -17 | -19 | -13 |  |  |  |  |
| 150 | 2 | 3 | 1 | 1 | 107 | 93 | 68 | 64 | 1 | -2 | -3 | 1 | 6 | 3 | 9 | 3 | -4 | -11 | -8 | -11 | 10 | -2 | 0 | -3 | -35 | -33 | -30 | -27 | 18 | 18 | 37 | 18 | 0.07 | 0 | 0.26 | 0.21 | -9 | -7 |  |  |  |  |  |  |
| 151 | 0 | 1 | 1 | 2 | -21 | -8 | -7 | -6 | 10 | 36 | 47 | 51 | 2 | 9 | -5 | -3 | -3 | -4 | -3 | -5 | 5 | 6 | 9 | 9 | 7 | 11 | 14 | 13 | 0 | 0 | -1 | -8 | 0.1 | 0.04 | -0.055 | 0.095 | -8 | -3 | -2 | -2 |  |  |  |  |
| 152 | -4 | -9 | -7 | -8 | -5 | 0 | -3 | 3 | -9 | -11 | -10 | -9 | -2 | -3 | -1 | 0 | -9 | -11 | -9 | -9 | 30 | 15 | -31 | -35 | 17 | 2 | 9 | 4 | -3 | -5 | -4 | -7 | 0.06 | 0.05 | 0.07 | 0.05 | -9 | -3 | -4 | -1 |  |  |  |  |
| 153 | -6 | -5 | -3 | -3 | 2 | 0 | -3 | 0 | -2 | -2 | -3 | -1 | 3 | 2 | -1 | 2 | 14 | 11 | 2 | -4 | -16 | -14 | -10 | -13 | 33 | 27 | 9 | 9 | -7 | 18 | 11 | 1 | 0.05 | 0.07 | 0.04 | 0.05 | -8 | -5 | -3 | -3 |  |  |  |  |
| 155 | -14 | -15 | -15 | -14 | -46 | -49 | -32 | -19 | -27 | -24 | -19 | -16 | 4 | 1 | 2 | 4 | 2 | 3 | 3 | 1 | -35 | -42 | -37 | -36 | 5 | 7 | 2 | 1 | -1 | -3 | 1 | 0 | 0.02 | 0 | 0.03 | 0.01 | -3 | -2 | -5 | -2 |  |  |  |  |
| 157 | 8 | 5 | 0 | 4 | 2 | -38 | -36 | -38 | 17 | 17 | 24 | 31 | -5 | -16 | -9 | 156 | 66 | 11 | 2 | 251 |  |  |  |  | 627 | -28 | -11 | 209 | -7 | -20 | -22 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 158 | 0 | 1 | 1 | 1 | -3 | -2 | -1 | 0 | 1 | 2 | 1 | 2 | 4 | 2 | 2 | 2 | -14 | -29 | -22 | -20 | 1 | -1 | 0 | -1 | -55 | -46 | -30 | -25 | -3 | -4 | -5 | -5 | 0.04 | 0.04 | 0.04 | 0.04 | -2 | 0 | -3 | 0 |  |  |  |  |
| 163 | 0 | 1 | 0 | -1 | 10 | 9 | 13 | 3 | -2 | -2 | -1 | -1 | 5 | 4 | 2 | 4 | 14 | 17 | 16 | 11 | -8 | -8 | -6 | -7 | 30 | 31 | 27 | 24 | -1 | 0 | -2 | -1 | 0.06 | 0.07 | 0.07 | 0.06 | -4 | -2 | -5 | 0 |  |  |  |  |
| 164 | -9 | -12 | -8 | -9 | 145 | 194 | 24 | 18 | -9 | -5 | -3 | -3 | -14 | -9 | -12 | -10 | -49 | -31 | -32 | -30 | -22 | -25 | -17 | -19 | -61 | -58 | -53 | -52 | -17 | -20 | -22 | -24 | 0.06 | 0.03 | 0.07 | 0.03 | -5 | -3 | -12 | -2 |  |  |  |  |
| 165 | 1 | 2 | 2 | 1 | 8 | -4 | -1 | 1 | 0 | 1 | 1 | 4 | 19 | 8 | 17 | 9 | 16 | 14 | 17 | 11 | 29 | 26 | 22 | 14 | 50 | 13 | 3 | -8 | -25 | -16 | 8 | 49 | 0.09 | 0 | 0 | 0.03 | -7 | -2 | -2 | 2 |  |  |  |  |
| 166 | 1 | 2 | 1 | 0 | 2 | 0 | -3 | 3 | -2 | -2 | 0 | -3 | 3 | -1 | 4 | 2 | -3 | -7 | -5 | -4 | -3 | -2 | -3 | -6 | -6 | -4 | -6 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 3 | 5 | 2 | 6 |  |  |  |  |
| 167 | 10 | 10 | 7 | 7 | -1 | 1 | -3 | 0 | 0 | -1 | 0 | 2 | -20 | -18 | -9 | -9 | -7 | -14 | -6 | -11 | 14 | 10 | 4 | 4 | -16 | -16 | 21 | 3 | -13 | -20 | -15 | -13 |  |  |  |  |  |  |  |  |  |  |  |  |

|  |  | SO42- and NO3-between + 10 and 20% |  | NH4+ , NO3-, Cl-, Na+, Mg2+, Ca2+, K+ and cond: between + 15-25% |  | pH: between + 0,1-0,2 pH-units |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | SO42- and NO3- more than + 20% |  | NH4+ , NO3-, Cl-, Na+, Mg2+, Ca2+, K+ and cond: between + 15-25% |  | pH: more than + 0,2 pH- units |

Table 11:	Reported results for filter samples expressed as % deviation from expected value.

| Air and aerosols |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Absorbing solution |  |  |  | Impregnated filter |  |  |  | Impregnated filter |  |  |  | Absorbing solution |  |  |  | Impregnated filter |  |  |  |
|  | SO2-S |  |  |  | SO2-S |  |  |  | HNO3-N |  |  |  | NO2-N |  |  |  | NH3-N |  |  |  |
|  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  | % deviation form expected value |  |  |  |
|  | A1 | A3 | A4 | A5 | B1 | B2 | B4 | B5 | B1 | B2 | B4 | B5 | C1 | C2 | C3 | C4 | J1 | J2 | J3 | J6 |
| 3 |  |  |  |  | -9 | -11 | 0 | -2 | -1 | -2 | -1 | 0 | -19 | -14 | -15 | -8 | 12 | -5 | -4 | -9 |
| 4 |  |  |  |  | 0 | 2 | 12 | 0 | 6 | 2 | 1 | 1 |  |  |  |  | -1 | -3 | -5 | -3 |
| 5 |  |  |  |  | 9 | 15 | -9 | 4 | 0 | -2 | -2 | 0 |  |  |  |  | -5 | -7 | -5 | -5 |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  | 3 | 1 | 3 | 3 | 11 | 14 | 9 | 16 |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  | -8 | -8 | -9 | -6 | -23 | -25 | -24 | -22 |
| 12 |  |  |  |  |  |  |  |  |  |  |  |  | 38 | 26 | 24 | 12 |  |  |  |  |
| 13 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | -1 | 2 | -1 | 2 |
| 15 | -5 | -1 | -4 | -1 | -4 | -26 | 1 | -1 | 4 | -10 | -8 | -8 | 3 | -2 | -3 | -2 | 10 | 9 | 18 | -4 |
| 16 |  |  |  |  | 2 | 4 | 1 | 1 | 4 | -3 | -3 | -3 | 8 | 0 | 4 | -1 | 5 | 8 | 3 | 13 |
| 17 | 26 | -2 | 4 | -14 |  |  |  |  | 7 | 1 | 1 | 1 |  |  |  |  |  |  |  |  |
| 19 | -17 | -4 | 6 | -2 |  |  |  |  |  |  |  |  | 46 | 49 | 47 | 50 |  |  |  |  |
| 20 |  |  |  |  | -8 | 1 | 14 | 3 |  |  |  |  | 0 | 1 | 0 | 0 | 1 | -1 | -1 | 3 |
| 21 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 22 |  |  |  |  | 4 | -1 | -12 | 12 | 23 | 19 | 30 | 28 |  |  |  |  |  |  |  |  |
| 23 | -9 | -10 | -3 | -4 |  |  |  |  |  |  |  |  | 5 | 1 | 0 | 3 |  |  |  |  |
| 31 |  |  |  |  | -23 | -19 | -12 | -7 | -8 | -14 | -10 | -4 | 0 | 2 | 1 | -1 | -6 | -13 | -18 | -6 |
| 32 |  |  |  |  | 2 | 3 | 4 | 18 | 11 | 7 | 11 | 8 | 8 | 6 | 4 | 5 | 2 | -2 | 0 | 5 |
| 33 |  |  |  |  | 23 | 5 | -80 | 6 | -17 | 3 | -15 | -4 | 11 | 853 | -6 | -7 | -5 | -1 | -3 | -6 |
| 34 | -5 | -11 | -12 | -11 | -13 | -6 | 5 | -79 | -11 | -9 | -77 | -78 | -24 | -9 | -16 | -4 | 2 | -14 | -19 | 6 |
| 35 |  |  |  |  |  |  |  |  |  |  |  |  | 0 | 4 | 0 | 2 |  |  |  |  |
| 36 | 7 | 6 | 5 | 9 | 5 | 2 | 9 | -3 | 5 | -10 | -7 | -8 | -3 | 1 | -1 | 3 | -6 | -6 | -4 | -6 |
| 38 |  |  |  |  | 10 | 14 | -6 | 11 |  |  |  |  | -3 | -1 | 0 | -3 |  |  |  |  |
| 39 |  |  |  |  | -2 | -4 | 3 | -7 | 18 | 12 | 9 | 7 | 32 | 15 | 10 | 5 | -8 | -10 | -9 | -8 |
| 41 |  |  |  |  | 12 | 16 | -17 | 20 | 0 | -1 | -2 | 3 |  |  |  |  | 44 | 30 | 20 | 53 |
| 45 |  |  |  |  |  |  |  |  |  |  |  |  | -1 | 4 | 15 | 13 | -1 | 1 | -1 | 3 |
| 116 |  |  |  |  | -25 | -27 | 3 | -13 | -26 | -18 | -18 | -14 |  |  |  |  | 4 | 4 | 6 | 1 |
| 158 |  |  |  |  | 12 | 20 | -11 | 3 | 2 | -1 | 2 | 3 |  |  |  |  | -4 | -10 | -9 | -7 |
| 163 |  |  |  |  | -13 | -18 | -17 | -9 | -1 | -7 | -7 | -4 |  |  |  |  |  |  |  |  |
| 164 |  |  |  |  | -46 | -70 | -19 | -12 | 5 | -1 | 1 | 2 |  |  |  |  |  |  |  |  |
| 165 |  |  |  |  | -65 | -87 | -19 | -13 | 1 | -6 | -6 | -1 |  |  |  |  | 25 | 22 | 20 | 18 |
| 173 | -100 | -100 | -100 | -100 |  |  |  |  |  |  |  |  | -3 | -5 | -3 | -5 |  |  |  |  |

|  |  |  | Results between 10 and 20% or between -10 and -20% from expected value |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | more than + 20% from expected value |  |  |  |  |  |  |  |

Table 12:	Relative random and systematic errors obtained by the different laboratories in the analysis of each parameter in the precipitation samples.

| Lab. no. | SO4 2- |  | NO3- |  | NH4+ |  | Mg2+ |  | H+ calc |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Random error % | Systematic error % | Random error % | Systematic error % | Random error % | Systematic error % | Random error % | Systematic error % | Random error % | Systematic error % |
| 1 | 1 | 1 | 1 | 0 | 1 | 4 | 3 | -12 | 3 | -6 |
| 2 | 1 | 8 | 3 | 9 | 2 | 8 | 1 | 17 | 10 | 65 |
| 3 | 0 | 1 | 1 | 0 | 3 | 23 | 1 | 1 | 2 | 4 |
| 4 | 1 | 0 | 1 | -3 | 1 | 0 | 2 | -16 | 4 | -7 |
| 5 | 0 | 0 | 1 | 1 | 1 | 0 | 2 | 1 | 3 | 1 |
| 7 | 1 | -2 | 1 | -1 | 2 | 1 | 3 | -1 | 2 | -9 |
| 8 | 0 | 4 | 2 | 5 | 1 | -2 | 1 | 4 | 4 | -8 |
| 10 | 2 | -4 | 1 | -3 | 2 | 0 | 2 | 6 | 12 | -15 |
| 12 | 2 | -5 | 1 | -4 | 2 | 4 | 4 | -3 | 3 | -3 |
| 13 | 2 | 9 | 5 | 10 | 2 | -1 | 3 | -3 | 9 | -3 |
| 14 | 0 | -2 | 1 | 0 | 3 | 1 | 5 | -8 | 6 | -23 |
| 15 | 2 | 3 | 2 | -3 | 4 | 0 | 8 | 20 | 6 | -8 |
| 16 | 1 | -2 | 1 | -4 | 3 | 3 | 4 | 1 | 4 | -3 |
| 17 | 7 | -18 | 2 | -13 | 3 | 14 | 3 | -26 | 9 | 18 |
| 19 | 4 | -7 | 1 | -9 | 5 | -7 | 1 | 0 | 9 | -25 |
| 20 | 1 | -2 | 1 | -1 | 1 | 7 | 14 | 22 | 6 | -3 |
| 21 | 1 | 2 | 1 | 0 | 1 | 0 | 1 | -4 | 4 | -6 |
| 22 | 1 | 2 | 7 | 17 | 3 | -4 | 2 | 3 | 3 | -10 |
| 23 | 2 | -3 | 1 | 1 | 3 | 9 | 4 | 12 | 12 | -13 |
| 24 | 1 | -6 | 2 | -10 | 5 | 0 | 3 | 8 | 231 | 6940 |
| 26 | 2 | -2 | 1 | 2 | 1 | 1 | 1 | -2 | 5 | -16 |
| 27 | 1 | 5 | 2 | 3 | 1 | 0 | 1 | 2 | 7 | -11 |
| 30 | 1 | -1 | 1 | -10 | 1 | -3 | 1 | -3 | 13 | -20 |
| 31 | 1 | -2 | 1 | -2 | 1 | 18 | 2 | 0 | 7 | -13 |
| 32 | 1 | -1 | 1 | 0 | 2 | -2 |  |  | 3 | -4 |
| 33 | 0 | -3 | 2 | 1 | 2 | 3 | 4 | -1 | 4 | -12 |
| 34 | 2 | -9 | 2 | -13 | 10 | 6 | 5 | -4 | 13 | -17 |
| 35 | 1 | 0 | 1 | -2 | 5 | 13 | 4 | 16 | 6 | -12 |
| 36 | 1 | -2 | 1 | -5 | 1 | -8 | 3 | -4 | 5 | -7 |
| 38 | 4 | 6 | 2 | 2 | 1 | 1 | 6 | -12 | 10 | 34 |
| 39 | 5 | 10 | 2 | 11 | 3 | 9 | 1 | 0 | 8 | -15 |
| 40 |  |  | 10 | 23 | 3 | 6 | 217 | 204 | 12 | 62 |
| 41 | 2 | 0 | 1 | -3 | 5 | 23 | 5 | -9 |  |  |
| 42 | 9 | 35 | 12 | 47 | 14 | 33 | 4 | 4 |  |  |
| 43 | 2 | -7 | 17 | -83 | 13 | -24 |  |  |  |  |
| 45 | 1 | -3 | 4 | -1 | 7 | 0 | 2 | -1 | 1 | -2 |
| 46 |  |  |  |  |  |  |  |  |  |  |
| 104 | 7 | 13 | 1 | -1 | 5 | -17 | 6 | -4 | 46 | -16 |
| 107 | 1 | -4 | 1 | -7 | 1 | -1 | 4 | 17 |  |  |
| 108 | 2 | -2 | 2 | -1 | 2 | -8 | 2 | -22 | 9 | -21 |
| 109 | 3 | -15 | 1 | -2 | 9 | 13 | 3 | -3 | 14 | -35 |
| 110 | 4 | 3 | 4 | -1 | 3 | -17 | 8 | -16 | 12 | -29 |
| 112 | 0 | 1 | 1 | 7 | 2 | -3 | 1 | -7 | 9 | -23 |
| 113 | 1 | -2 | 1 | 3 | 19 | 6 | 3 | -8 | 5 | -15 |
| 114 | 2 | -2 | 1 | 0 | 2 | -2 | 6 | 17 | 5 | -3 |
| 115 | 1 | 4 | 2 | 4 | 4 | 0 | 4 | -12 | 10 | -24 |
| 116 | 2 | -1 | 2 | -2 | 3 | 6 | 1 | -4 | 18 | 5 |
| 118 | 6 | -2 | 1 | 1 | 0 | -8 | 2 | -7 | 2 | 2 |
| 120 | 2 | -3 | 1 | -2 | 7 | -35 | 2 | -4 | 2 | 0 |
| 121 | 2 | -5 | 72 | 331 | 10 | 27 | 1 | -3 | 5 | -6 |
| 124 | 3 | -2 | 4 | 2 | 10 | 19 | 8 | 4 |  |  |
| 125 | 1 | -1 | 1 | -1 | 8 | 8 | 4 | 8 |  |  |
| 126 | 5 | -10 | 3 | -4 | 14 | -34 | 16 | -19 | 6 | -12 |
| 133 | 32 | -19 | 7 | -14 | 18 | 26 | 2 | -3 | 19 | -43 |
| 140 | 2 | -2 | 2 | -2 | 2 | 4 | 3 | -3 | 8 | -16 |
| 145 | 15 | -12 | 6 | -29 | 18 | 8 | 50 | 28 |  |  |
| 150 | 1 | 2 | 2 | 0 | 10 | 76 | 2 | -7 | 12 | -26 |
| 151 | 1 | 1 | 23 | 41 | 3 | -9 | 1 | -4 | 17 | -13 |
| 152 | 2 | -6 | 2 | -10 | 3 | -2 | 3 | -8 | 5 | -11 |
| 153 | 2 | -4 | 1 | -2 | 2 | 0 | 11 | 4 | 5 | -12 |
| 155 | 3 | -14 | 1 | -21 | 5 | -32 | 1 | 2 | 3 | -3 |
| 157 | 4 | 4 | 11 | 21 | 23 | -35 | 103 | 55 | 33 | -93 |
| 158 | 0 | 0 | 1 | 1 | 1 | -1 | 2 | -20 | 4 | -8 |
| 163 | 1 | 0 | 0 | -1 | 5 | 7 | 4 | 12 | 4 | -13 |
| 164 | 3 | -10 | 2 | -4 | 46 | 62 | 23 | -28 | 6 | -8 |
| 165 | 0 | 2 | 2 | 1 | 3 | 0 | 7 | 12 | 12 | -3 |
| 166 | 1 | 1 | 1 | -2 | 3 | 1 | 1 | -4 | 2 | 0 |
| 167 | 4 | 8 | 1 | 0 | 2 | -1 | 3 | -10 |  |  |

Table 12, cont.

| Lab. no. | Na+ |  | Cl- |  | K+ |  | Cond. |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Random error % | Systematic error % | Random error % | Systematic error % | Random error % | Systematic error % | Random error % | Systematic error % |
| 1 | 1 | -3 | 1 | -2 | 2 | -8 | 4 | -2 |
| 2 | 3 | -4 | 4 | 26 | 5 | 18 | 2 | -1 |
| 3 | 2 | -5 | 1 | -2 | 2 | 0 | 2 | 1 |
| 4 | 1 | 5 |  |  | 4 | -9 | 3 | 3 |
| 5 | 1 | 3 | 1 | 1 | 3 | 5 | 1 | 6 |
| 7 | 2 | 6 | 4 | 5 | 1 | 2 | 6 | -11 |
| 8 | 0 | 0 | 3 | 1 | 3 | -5 | 2 | -3 |
| 10 | 2 | -11 | 4 | -11 | 3 | 15 | 1 | 1 |
| 12 | 0 | -4 | 1 | -7 | 2 | -7 | 2 | 2 |
| 13 | 1 | 2 | 6 | 11 | 2 | -4 | 12 | 2 |
| 14 | 1 | 1 | 6 | -19 | 3 | -10 | 2 | 6 |
| 15 | 1 | -3 | 6 | -7 | 2 | -2 | 2 | 2 |
| 16 | 3 | -3 | 2 | -5 | 3 | -4 | 2 | -2 |
| 17 |  |  | 40 | 104 |  |  | 4 | 0 |
| 19 | 2 | 3 |  |  | 3 | 6 | 2 | 1 |
| 20 | 3 | 11 | 1 | -1 | 3 | 5 | 19 | 1 |
| 21 | 2 | 1 | 1 | -5 | 1 | -3 | 1 | 3 |
| 22 | 7 | -11 | 9 | 7 | 11 | -17 | 4 | -20 |
| 23 | 2 | 0 | 1 | -3 | 5 | 0 | 4 | -4 |
| 24 | 5 | -2 | 13 | -68 | 2 | -6 | 3 | 6 |
| 26 | 0 | 0 | 2 | 1 | 1 | -2 |  |  |
| 27 | 0 | 1 | 1 | 1 | 0 | -2 | 2 | -1 |
| 30 | 1 | 4 | 5 | 0 | 2 | -8 | 4 | 0 |
| 31 | 2 | -7 | 2 | -3 | 4 | -6 | 2 | 1 |
| 32 | 2 | -1 | 2 | 3 | 2 | 0 | 2 | -2 |
| 33 | 3 | -4 | 3 | -1 | 3 | -5 | 2 | 0 |
| 34 |  |  | 16 | -61 |  |  | 4 | -6 |
| 35 | 4 | -11 | 3 | -8 |  |  | 2 | -5 |
| 36 | 1 | -2 | 3 | 10 | 4 | -8 | 3 | -9 |
| 38 | 6 | -34 | 5 | 2 | 5 | -22 | 2 | 2 |
| 39 | 2 | 8 | 2 | 11 | 0 | 0 | 3 | 0 |
| 40 | 20 | 22 | 6 | 1233 | 23 | 2 | 5 | -9 |
| 41 | 3 | 7 | 7 | -19 | 2 | 8 |  |  |
| 42 | 4 | -16 | 15 | 41 | 6 | -9 |  |  |
| 43 |  |  | 61 | 87 |  |  |  |  |
| 45 | 2 | 2 | 3 | -6 | 3 | -3 | 7 | -4 |
| 46 |  |  |  |  |  |  | 1 | 0 |
| 104 | 4 | -14 | 6 | 21 | 4 | -10 | 20 | -8 |
| 107 | 1 | -4 | 6 | -17 | 1 | -7 | 3 | -3 |
| 108 | 4 | 2 | 22 | 88 | 2 | 4 | 4 | -9 |
| 109 | 15 | -27 |  |  |  |  | 1 | -3 |
| 110 | 2 | 5 |  |  |  |  | 4 | -6 |
| 112 | 0 | 4 | 10 | 12 | 1 | -4 | 3 | -7 |
| 113 | 3 | 0 | 14 | 0 | 19 | 13 | 4 | -9 |
| 114 | 0 | -1 | 6 | -7 | 9 | -9 | 3 | 0 |
| 115 | 3 | -9 | 9 | 10 | 7 | -27 | 1 | -4 |
| 116 | 2 | 1 | 4 | -12 | 0 | 3 | 11 | 4 |
| 118 | 4 | -18 | 3 | -14 | 7 | -17 | 3 | -4 |
| 120 | 0 | 2 | 26 | -5 | 9 | -7 | 2 | 2 |
| 121 | 2 | -67 | 8 | -6 |  |  | 3 | 4 |
| 124 | 1 | 6 | 2 | -1 | 6 | -3 | 2 | -5 |
| 125 | 1 | -1 | 4 | 0 | 3 | 1 |  |  |
| 126 | 1 | 5 | 3 | -16 | 2 | -8 | 2 | 8 |
| 133 | 7 | -39 | 11 | 47 | 3 | 1 | 7 | -20 |
| 140 | 1 | 0 | 3 | -7 | 10 | 22 | 4 | -11 |
| 145 | 6 | 3 | 11 | 5 |  |  | 3 | -16 |
| 150 | 3 | 4 | 4 | -1 | 9 | 21 |  |  |
| 151 | 6 | -1 | 4 | 8 | 3 | 0 | 4 | -3 |
| 152 | 1 | -1 | 27 | -14 | 1 | -4 | 4 | -3 |
| 153 | 1 | 2 | 3 | -12 | 12 | 5 | 3 | -4 |
| 155 | 1 | 3 | 10 | -39 | 1 | -1 | 1 | -3 |
| 157 | 88 | -7 |  |  | 591 | -14 |  |  |
| 158 | 0 | 3 | 1 | 0 | 0 | -4 | 1 | -1 |
| 163 | 1 | 3 | 2 | -6 | 1 | -1 | 2 | -3 |
| 164 | 3 | -11 | 4 | -20 | 2 | -20 | 3 | -5 |
| 165 | 5 | 13 | 4 | 20 | 27 | -5 | 5 | -2 |
| 166 | 2 | 2 | 3 | -3 |  |  | 1 | 4 |

Table 13:	Analytical results for sulphur dioxide in absorbing solution.

SO2-S in absorbing solution                 SO2-S in absorbing solution

Sample no.: A1                              Sample no.: A3

Theoretical value:        0.100             Theoretical value:        0.220

Unit: µg S/ml                               Unit: µg S/ml

Run 1:                                      Run 1:

Number of laboratories:       7             Number of laboratories:       7

Arithmetic mean value:    0.085             Arithmetic mean value:    0.182

Median:                   0.095             Median:                   0.211

Standard deviation        0.040             Standard deviation        0.081

Rel. st. deviation (%)   46.973             Rel. st. deviation (%)   44.653

Run 2:                                      Run 2:

Number of laboratories:       6             Number of laboratories:       6

Arithmetic mean value:    0.099             Arithmetic mean value:    0.212

Median:                   0.095             Median:                   0.213

Standard deviation        0.015             Standard deviation        0.014

Rel. st. deviation (%)   15.199             Rel. st. deviation (%)    6.605

Results in decreasing order:                Results in decreasing order:

17    0.126       23    0.091               36    0.233       23    0.198

36    0.107       19    0.083               15    0.218       34    0.195

15    0.095      173    0.000 (*)           17    0.216      173    0.000 (*)

34    0.095                                 19    0.211

SO2-S in absorbing solution                 SO2-S in absorbing solution

Sample no.: A4                              Sample no.: A5

Theoretical value:        0.160             Theoretical value:        0.320

Unit: µg S/ml                               Unit: µg S/ml

Run 1:                                      Run 1:

Number of laboratories:       7             Number of laboratories:       7

Arithmetic mean value:    0.136             Arithmetic mean value:    0.264

Median:                   0.154             Median:                   0.307

Standard deviation        0.061             Standard deviation        0.119

Rel. st. deviation (%)   44.740             Rel. st. deviation (%)   45.007

Run 2:                                      Run 2:

Number of laboratories:       6             Number of laboratories:       6

Arithmetic mean value:    0.159             Arithmetic mean value:    0.308

Median:                   0.160             Median:                   0.310

Standard deviation        0.011             Standard deviation        0.026

Rel. st. deviation (%)    7.104             Rel. st. deviation (%)    8.462

Results in decreasing order:                Results in decreasing order:

19    0.170       15    0.153               36    0.350       34    0.286

36    0.168       34    0.141               15    0.318       17    0.276

17    0.166      173    0.000 (*)           19    0.314      173    0.000 (*)

23    0.154                                 23    0.307

Table 14:	Analytical results for sulphur dioxide on impregnated filter.

SO2-S on impregnated filter                 SO2-S on impregnated filter

Sample no.: B1                              Sample no.: B2

Theoretical value:       24.050             Theoretical value:       16.000

Unit: ug S/filter                           Unit: ug S/filter

Run 1:                                      Run 1:

Number of laboratories:      20             Number of laboratories:      20

Arithmetic mean value:   22.484             Arithmetic mean value:   14.519

Median:                  23.740             Median:                  15.995

Standard deviation        5.012             Standard deviation        4.349

Rel. st. deviation (%)   22.289             Rel. st. deviation (%)   29.956

Run 2:                                      Run 2:

Number of laboratories:      19             Number of laboratories:      18

Arithmetic mean value:   23.231             Arithmetic mean value:   15.744

Median:                  23.980             Median:                  16.260

Standard deviation        3.840             Standard deviation        2.239

Rel. st. deviation (%)   16.530             Rel. st. deviation (%)   14.223

Results in decreasing order:                Results in decreasing order:

33   29.596       39   23.500              158   19.150       22   15.770

158   26.920       15   23.100               41   18.600       39   15.400

41   26.850       20   22.240                5   18.370       34   15.041

38   26.400        3   21.849               38   18.200        3   14.299

5   26.170      163   20.980               33   16.871      163   13.180

36   25.320       34   20.831               16   16.680       31   13.014

22   25.050       31   18.504               32   16.480       15   11.800

32   24.610      116   17.940               36   16.360      116   11.660

16   24.490      164   13.050                4   16.300      164    4.880 (*)

4   23.980      165    8.300 (*)           20   16.220      165    2.100 (*)

SO2-S on impregnated filter                 SO2-S on impregnated filter

Sample no.: B4                              Sample no.: B5

Theoretical value:       60.090             Theoretical value:       76.110

Unit: ug S/filter                           Unit: ug S/filter

Run 1:                                      Run 1:

Number of laboratories:      20             Number of laboratories:      20

Arithmetic mean value:   55.578             Arithmetic mean value:   73.431

Median:                  58.245             Median:                  75.745

Standard deviation       12.000             Standard deviation       15.419

Rel. st. deviation (%)   21.592             Rel. st. deviation (%)   20.998

Run 2:                                      Run 2:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:   57.871             Arithmetic mean value:   76.469

Median:                  60.190             Median:                  76.290

Standard deviation        6.405             Standard deviation        7.495

Rel. st. deviation (%)   11.068             Rel. st. deviation (%)    9.801

Results in decreasing order:                Results in decreasing order:

20   68.240       38   56.300               41   91.300       15   75.200

4   67.600        5   54.900               32   89.950        3   74.299

36   65.500      158   53.280               22   85.020       36   73.500

34   62.810       22   52.804               38   84.700       31   70.924

32   62.272       31   52.804               33   80.397       39   70.900

116   62.180      163   49.950                5   79.110      163   68.980

39   62.000       41   49.730               20   78.340      164   66.670

15   60.900      164   48.800              158   78.310      165   66.400

16   60.480      165   48.800               16   76.580      116   66.040

3   60.190       33   12.014 (*)            4   76.290       34   15.716 (*)

Table 15:	Analytical results for nitric acid on impregnated filter.

HNO3-N on impregnated filter                HNO3-N on impregnated filter

Sample no.: B1                              Sample no.: B2

Theoretical value:       11.450             Theoretical value:       16.360

Unit: ug N/filter                           Unit: ug N/filter

Run 1:                                      Run 1:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:   11.593             Arithmetic mean value:   16.021

Median:                  11.720             Median:                  16.060

Standard deviation        1.277             Standard deviation        1.460

Rel. st. deviation (%)   11.020             Rel. st. deviation (%)    9.116

Run 2:                                      Run 2:

Number of laboratories:      18             Number of laboratories:      18

Arithmetic mean value:   11.764             Arithmetic mean value:   15.825

Median:                  11.810             Median:                  16.057

Standard deviation        1.067             Standard deviation        1.218

Rel. st. deviation (%)    9.068             Rel. st. deviation (%)    7.700

Results in decreasing order:                Results in decreasing order:

22   14.050      165   11.600               22   19.550 (*)    3   16.054

39   13.500        5   11.460               39   18.400       16   15.870

32   12.700       41   11.410               32   17.510      165   15.300

20   12.280        3   11.354               33   16.810      163   15.250

4   12.190      163   11.350                4   16.680       34   14.833

36   12.060       31   10.554               20   16.540       36   14.770

164   11.980       34   10.185              164   16.240       15   14.700

16   11.940       33    9.516               41   16.210       31   14.114

15   11.900      116    8.510 (*)          158   16.150      116   13.350

158   11.720                                  5   16.060

HNO3-N on impregnated filter                HNO3-N on impregnated filter

Sample no.: B4                              Sample no.: B5

Theoretical value:       22.900             Theoretical value:       26.180

Unit: ug N/filter                           Unit: ug N/filter

Run 1:                                      Run 1:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:   21.676             Arithmetic mean value:   25.174

Median:                  22.450             Median:                  26.130

Standard deviation        4.656             Standard deviation        5.201

Rel. st. deviation (%)   21.481             Rel. st. deviation (%)   20.660

Run 2:                                      Run 2:

Number of laboratories:      18             Number of laboratories:      18

Arithmetic mean value:   22.591             Arithmetic mean value:   26.250

Median:                  22.500             Median:                  26.172

Standard deviation        2.468             Standard deviation        2.316

Rel. st. deviation (%)   10.923             Rel. st. deviation (%)    8.824

Results in decreasing order:                Results in decreasing order:

22   29.850       16   22.280               22   33.570      165   25.800

32   25.370      165   21.500               32   28.340       16   25.360

39   24.900       36   21.300               39   27.900       31   25.154

158   23.330      163   21.250              158   27.030      163   25.050

164   23.220       15   21.100               41   26.910       33   25.005

20   23.160       31   20.574              164   26.640       15   24.200

4   23.030       33   19.425                4   26.480       36   23.970

3   22.664      116   18.690               20   26.320      116   22.420

5   22.550       34    5.195 (*)            3   26.214       34    5.812 (*)

41   22.450                                  5   26.130

Table 16:	Analytical results for nitrogen dioxide in absorbing solution.

NO2-N in absorbing solution                 NO2-N in absorbing solution

Sample no.: C1                              Sample no.: C2

Theoretical value:        0.037             Theoretical value:        0.085

Unit: ug N/ml                               Unit: ug N/ml

Run 1:                                      Run 1:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:    0.039             Arithmetic mean value:    0.126

Median:                   0.037             Median:                   0.086

Standard deviation        0.006             Standard deviation        0.166

Rel. st. deviation (%)   16.607             Rel. st. deviation (%)  131.343

Run 2:                                      Run 2:

Number of laboratories:      18             Number of laboratories:      18

Arithmetic mean value:    0.038             Arithmetic mean value:    0.088

Median:                   0.037             Median:                   0.086

Standard deviation        0.005             Standard deviation        0.012

Rel. st. deviation (%)   14.351             Rel. st. deviation (%)   13.843

Results in decreasing order:                Results in decreasing order:

19    0.054 (*)   31    0.037               33    0.810 (*)   23    0.086

12    0.051       20    0.037               19    0.127       20    0.086

39    0.049       35    0.037               12    0.107       16    0.085

33    0.041       36    0.036               39    0.098       38    0.084

32    0.040       38    0.036               32    0.090       15    0.083

16    0.040      173    0.036               45    0.089      173    0.081

23    0.039       10    0.034               35    0.088       10    0.078

15    0.038        3    0.030               31    0.087       34    0.077

8    0.038       34    0.028               36    0.086        3    0.073

45    0.037                                  8    0.086

NO2-N in absorbing solution                 NO2-N in absorbing solution

Sample no.: C3                              Sample no.: C4

Theoretical value:        0.068             Theoretical value:        0.119

Unit: ug N/ml                               Unit: ug N/ml

Run 1:                                      Run 1:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:    0.070             Arithmetic mean value:    0.123

Median:                   0.068             Median:                   0.119

Standard deviation        0.010             Standard deviation        0.015

Rel. st. deviation (%)   13.713             Rel. st. deviation (%)   12.344

Run 2:                                      Run 2:

Number of laboratories:      18             Number of laboratories:      18

Arithmetic mean value:    0.068             Arithmetic mean value:    0.120

Median:                   0.068             Median:                   0.118

Standard deviation        0.006             Standard deviation        0.007

Rel. st. deviation (%)    9.441             Rel. st. deviation (%)    5.733

Results in decreasing order:                Results in decreasing order:

19    0.100 (*)   23    0.068               19    0.179 (*)   31    0.118

12    0.084       38    0.068               45    0.134       16    0.118

45    0.078       36    0.067               12    0.133       15    0.117

39    0.075       15    0.066               39    0.125       38    0.116

32    0.071      173    0.066               32    0.125       34    0.114

16    0.071       33    0.064               23    0.123      173    0.113

8    0.070       10    0.062                8    0.123       10    0.112

31    0.069        3    0.058               36    0.122       33    0.111

35    0.068       34    0.057               35    0.121        3    0.110

20    0.068                                 20    0.119

Table 17:	Analytical results for ammonia on impregnated filter. The reported results are corrected for an average blank value (J3 and J5).

NH3-N on impregnated filter                 NH3-N on impregnated filter

Sample no.: J1                              Sample no.: J2

Theoretical value:       14.040             Theoretical value:       22.060

Unit: ug N/filter                           Unit: ug N/filter

Run 1:                                      Run 1:

Number of laboratories:      20             Number of laboratories:      20

Arithmetic mean value:   14.426             Arithmetic mean value:   21.986

Median:                  14.082             Median:                  21.684

Standard deviation        1.920             Standard deviation        2.795

Rel. st. deviation (%)   13.310             Rel. st. deviation (%)   12.712

Run 2:                                      Run 2:

Number of laboratories:      19             Number of laboratories:      19

Arithmetic mean value:   14.119             Arithmetic mean value:   21.639

Median:                  13.930             Median:                  21.610

Standard deviation        1.379             Standard deviation        2.390

Rel. st. deviation (%)    9.767             Rel. st. deviation (%)   11.043

Results in decreasing order:                Results in decreasing order:

41   20.260 (*)    4   13.930               41   28.570 (*)   32   21.610

165   17.500       45   13.850              165   27.000        4   21.500

3   15.755       13   13.840                8   25.050        3   21.035

8   15.550      158   13.540               15   24.035       36   20.830

15   15.435        5   13.380               16   23.890        5   20.470

16   14.700       33   13.312              116   22.925      158   19.850

116   14.635       31   13.161               13   22.600       39   19.748

32   14.260       36   13.130               45   22.250       31   19.131

34   14.255       39   12.948               20   21.935       34   19.005

20   14.235       10   10.848               33   21.759       10   16.527

NH3-N on impregnated filter                 NH3-N on impregnated filter

Sample no.: J3                              Sample no.: J6

Theoretical value:       28.070             Theoretical value:       11.030

Unit: ug N/filter                           Unit: ug N/filter

Run 1:                                      Run 1:

Number of laboratories:      20             Number of laboratories:      20

Arithmetic mean value:   27.695             Arithmetic mean value:   11.273

Median:                  27.417             Median:                  10.928

Standard deviation        3.344             Standard deviation        1.673

Rel. st. deviation (%)   12.075             Rel. st. deviation (%)   14.839

Run 2:                                      Run 2:

Number of laboratories:      20             Number of laboratories:      19

Arithmetic mean value:   27.695             Arithmetic mean value:   10.977

Median:                  27.417             Median:                  10.720

Standard deviation        3.344             Standard deviation        1.050

Rel. st. deviation (%)   12.075             Rel. st. deviation (%)    9.563

Results in decreasing order:                Results in decreasing order:

165   33.700       33   27.183               41   16.900 (*)    4   10.720

41   33.550        3   26.925              165   13.000       15   10.595

15   33.025       36   26.920                8   12.750        5   10.460

8   30.650        4   26.780               16   12.510       36   10.410

116   29.825        5   26.720               34   11.645       33   10.405

16   28.910       39   25.648               32   11.540       31   10.351

32   28.070      158   25.590               20   11.355      158   10.250

45   27.860       31   22.930               45   11.310       39   10.149

20   27.815       34   22.765               13   11.290        3   10.055

13   27.650       10   21.376              116   11.135       10    8.627

Table 18:	Analytical results for sulphate in precipitations samples.

Sulphate in precipitation                   Sulphate in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        1.402             Theoretical value:        1.030

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    1.371             Arithmetic mean value:    1.024

Median:                   1.383             Median:                   1.021

Standard deviation        0.143             Standard deviation        0.084

Rel. st. deviation (%)   10.460             Rel. st. deviation (%)    8.214

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      63

Arithmetic mean value:    1.375             Arithmetic mean value:    1.019

Median:                   1.383             Median:                   1.021

Standard deviation        0.081             Standard deviation        0.063

Rel. st. deviation (%)    5.912             Rel. st. deviation (%)    6.191

Results in decreasing order:                Results in decreasing order:

42    1.879 (*)   38    1.380               42    1.402 (*)   32    1.021

104    1.670 (*)  107    1.380              104    1.200 (*)  116    1.021

39    1.590       30    1.378               39    1.168      114    1.020

167    1.548      113    1.375              133    1.167       31    1.013

157    1.510       32    1.371              167    1.138       45    1.011

2    1.489       33    1.370               13    1.133       14    1.010

13    1.487       31    1.369                2    1.115      140    1.010

27    1.447       45    1.365               38    1.100      124    1.010

110    1.445       16    1.362               27    1.084       16    1.008

41    1.440       20    1.360              115    1.080       36    1.008

8    1.433       36    1.355              157    1.080       26    1.003

115    1.430      152    1.350               15    1.077       33    1.000

150    1.430      140    1.350                8    1.067      108    1.000

114    1.430       26    1.347              150    1.060       20    1.000

22    1.423       24    1.346               21    1.058       10    0.992

116    1.421      108    1.340               22    1.056       23    0.991

1    1.420      120    1.330              165    1.053      120    0.990

165    1.420       10    1.328              166    1.050      107    0.985

166    1.420       23    1.325                1    1.050      118    0.980

124    1.420      153    1.320              112    1.046      153    0.980

21    1.418      121    1.310                3    1.043      121    0.970

3    1.415       12    1.310              151    1.042       12    0.970

4    1.409       43    1.299              163    1.040       43    0.961

151    1.408      118    1.280              110    1.040       19    0.951

15    1.408       34    1.276                4    1.038       24    0.947

112    1.408      164    1.270               41    1.036      152    0.940

5    1.405       19    1.246              158    1.036       34    0.934

158    1.403      126    1.225                5    1.033      126    0.918

163    1.400      155    1.202              125    1.030      164    0.910

7    1.400      109    1.200               35    1.024      155    0.875

125    1.390       17    1.103              113    1.023      145    0.869

14    1.390      145    0.949 (*)           30    1.023      109    0.860

35    1.387      133    0.733 (*)            7    1.021       17    0.830 (*)

Sulphate in precipitation                   Sulphate in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.750             Theoretical value:        0.868

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    0.739             Arithmetic mean value:    0.859

Median:                   0.738             Median:                   0.855

Standard deviation        0.060             Standard deviation        0.070

Rel. st. deviation (%)    8.115             Rel. st. deviation (%)    8.109

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      63

Arithmetic mean value:    0.737             Arithmetic mean value:    0.859

Median:                   0.738             Median:                   0.859

Standard deviation        0.041             Standard deviation        0.047

Rel. st. deviation (%)    5.554             Rel. st. deviation (%)    5.519

Results in decreasing order:                Results in decreasing order:

42    1.001 (*)   26    0.735               42    1.198 (*)   26    0.851

13    0.871 (*)   16    0.733              104    0.970       20    0.850

104    0.850      116    0.733              110    0.968       14    0.850

2    0.815       36    0.731                2    0.952      140    0.850

39    0.806       30    0.731               13    0.945      108    0.850

38    0.800      153    0.730               39    0.938      116    0.849

115    0.800      120    0.730              167    0.932       31    0.847

167    0.800      125    0.730               38    0.930        7    0.846

8    0.787       14    0.730               27    0.915      113    0.846

27    0.786       31    0.728                8    0.903       36    0.845

15    0.778        7    0.726              115    0.900       23    0.842

110    0.771       10    0.724              157    0.900       16    0.841

118    0.770       23    0.723               15    0.896       33    0.840

165    0.767      114    0.720              151    0.889      120    0.840

112    0.761       33    0.720               22    0.884      114    0.840

151    0.760      124    0.720               21    0.883      153    0.840

166    0.760       45    0.718              150    0.880       10    0.834

150    0.760      113    0.711              112    0.880      107    0.825

21    0.758       12    0.710              165    0.879      124    0.820

1    0.757      107    0.707                1    0.878      121    0.820

3    0.755       24    0.700              158    0.874       12    0.820

158    0.754      152    0.700              118    0.870       43    0.815

157    0.750      121    0.700                3    0.870       24    0.811

163    0.748       19    0.697              166    0.870       19    0.810

35    0.748      164    0.690               41    0.869      126    0.804

4    0.747       34    0.674                5    0.868      152    0.800

5    0.743       43    0.669               32    0.867      164    0.790

41    0.743      145    0.669               35    0.867      145    0.786

32    0.741      126    0.664               45    0.866       34    0.782

20    0.740      155    0.635              125    0.861      155    0.747

22    0.740      109    0.630                4    0.860      109    0.730

140    0.740       17    0.616 (*)           30    0.859       17    0.706 (*)

108    0.740      133    0.566 (*)          163    0.859      133    0.666 (*)

Table 19:	Analytical results for nitrate in precipitations samples.

Nitrate in precipitation                    Nitrate in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.307             Theoretical value:        0.325

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      67             Number of laboratories:      67

Arithmetic mean value:    0.316             Arithmetic mean value:    0.337

Median:                   0.304             Median:                   0.320

Standard deviation        0.136             Standard deviation        0.144

Rel. st. deviation (%)   42.886             Rel. st. deviation (%)   42.845

Run 2:                                      Run 2:

Number of laboratories:      66             Number of laboratories:      65

Arithmetic mean value:    0.301             Arithmetic mean value:    0.325

Median:                   0.304             Median:                   0.320

Standard deviation        0.048             Standard deviation        0.037

Rel. st. deviation (%)   15.958             Rel. st. deviation (%)   11.300

Results in decreasing order:                Results in decreasing order:

121    1.340 (*)   32    0.304              121    1.430 (*)   20    0.320

42    0.451       45    0.303               42    0.488      109    0.320

151    0.399        7    0.302              151    0.443      108    0.320

40    0.377       15    0.301               22    0.386      163    0.320

157    0.360      153    0.300               40    0.382      150    0.320

39    0.347      166    0.300              157    0.380      153    0.320

112    0.337      108    0.300               39    0.375      116    0.319

13    0.330      104    0.300               13    0.356       31    0.317

22    0.329      163    0.300                2    0.356       10    0.316

2    0.329      120    0.300              112    0.355       35    0.316

115    0.328       35    0.299              115    0.348       15    0.316

113    0.326       31    0.297                8    0.343      110    0.316

26    0.320        4    0.296               26    0.337      140    0.312

124    0.320       41    0.292               27    0.334        4    0.312

38    0.320       10    0.292               45    0.331       12    0.310

8    0.318      140    0.291              104    0.330      120    0.310

27    0.316       16    0.291              158    0.330      126    0.310

5    0.312       36    0.291              113    0.330       41    0.310

116    0.312       12    0.290                5    0.330      164    0.310

14    0.310      109    0.290              118    0.330      124    0.310

20    0.310      126    0.290               38    0.330       36    0.308

158    0.310       24    0.284               33    0.330       16    0.305

118    0.310      164    0.280               14    0.330      107    0.302

33    0.310      107    0.280               23    0.328       19    0.296

150    0.310      152    0.280              165    0.328      152    0.290

114    0.307       19    0.272               32    0.327       24    0.288

165    0.307       30    0.269                3    0.327       30    0.288

1    0.306       34    0.266                1    0.325       34    0.281

23    0.306       17    0.260               21    0.325      133    0.277

125    0.306      133    0.248              167    0.323       17    0.276

167    0.306      155    0.225              114    0.321      155    0.247

3    0.306      145    0.157              125    0.321      145    0.236

110    0.305       43    0.052              166    0.320       43    0.042 (*)

21    0.304                                  7    0.320

Nitrate in precipitation                    Nitrate in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.442             Theoretical value:        0.505

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      67             Number of laboratories:      67

Arithmetic mean value:    0.461             Arithmetic mean value:    0.532

Median:                   0.440             Median:                   0.504

Standard deviation        0.197             Standard deviation        0.226

Rel. st. deviation (%)   42.692             Rel. st. deviation (%)   42.378

Run 2:                                      Run 2:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    0.438             Arithmetic mean value:    0.506

Median:                   0.440             Median:                   0.504

Standard deviation        0.068             Standard deviation        0.081

Rel. st. deviation (%)   15.499             Rel. st. deviation (%)   16.077

Results in decreasing order:                Results in decreasing order:

121    1.950 (*)  104    0.440              121    2.230 (*)   41    0.504

42    0.653        3    0.438              151    0.764       31    0.502

151    0.649      163    0.436               42    0.763        7    0.501

40    0.555       35    0.436               40    0.662       14    0.500

157    0.550      125    0.435              157    0.660      104    0.500

22    0.513       31    0.435               22    0.593      120    0.500

13    0.488        7    0.433               13    0.575      126    0.500

2    0.480       41    0.432               39    0.558      153    0.500

39    0.475      109    0.430                2    0.554      163    0.500

124    0.470      153    0.430                8    0.535      109    0.500

8    0.467       12    0.430               27    0.529       20    0.500

112    0.464      164    0.430              112    0.527      125    0.499

113    0.461      150    0.430              165    0.525      116    0.499

27    0.458      120    0.430              110    0.522        4    0.499

158    0.448        4    0.427              108    0.520       10    0.495

5    0.447      116    0.427               33    0.520       16    0.491

165    0.445      110    0.425               38    0.520      164    0.490

23    0.444       10    0.424              158    0.517      166    0.490

167    0.443       16    0.423              167    0.515       12    0.490

115    0.443       15    0.419                5    0.515       15    0.488

26    0.443       36    0.415               23    0.514       36    0.481

140    0.442      107    0.410              114    0.512      107    0.476

1    0.442      126    0.410              115    0.512       19    0.472

21    0.441       19    0.407                1    0.511       45    0.468

32    0.441      133    0.406              118    0.510       24    0.466

45    0.441       24    0.400              124    0.510       30    0.466

114    0.441      152    0.400              150    0.510      152    0.460

33    0.440       30    0.398               21    0.510       17    0.453

108    0.440       34    0.381              113    0.510       34    0.440

20    0.440       17    0.380               32    0.508      155    0.424

38    0.440      155    0.358               35    0.508      133    0.400

166    0.440      145    0.336                3    0.506      145    0.385

14    0.440       43    0.072              140    0.506       43    0.084

118    0.440                                 26    0.504

Table 20:	Analytical results for ammonium in precipitations sample.

Ammonium in precipitation                   Ammonium in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.147             Theoretical value:        0.160

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      67             Number of laboratories:      67

Arithmetic mean value:    0.155             Arithmetic mean value:    0.169

Median:                   0.150             Median:                   0.160

Standard deviation        0.040             Standard deviation        0.049

Rel. st. deviation (%)   25.953             Rel. st. deviation (%)   28.874

Run 2:                                      Run 2:

Number of laboratories:      65             Number of laboratories:      64

Arithmetic mean value:    0.149             Arithmetic mean value:    0.164

Median:                   0.150             Median:                   0.160

Standard deviation        0.025             Standard deviation        0.024

Rel. st. deviation (%)   16.709             Rel. st. deviation (%)   14.506

Results in decreasing order:                Results in decreasing order:

164    0.360 (*)  166    0.150              164    0.470 (*)  153    0.160

150    0.304 (*)  153    0.150              150    0.309 (*)   27    0.160

133    0.217      157    0.150              133    0.225      166    0.160

109    0.200       22    0.148               42    0.218       13    0.160

3    0.191       34    0.148                3    0.209        5    0.159

42    0.190       38    0.148              121    0.207        7    0.159

121    0.190      167    0.145               41    0.203      114    0.158

31    0.185       10    0.145               31    0.202       30    0.158

41    0.184       27    0.145              124    0.190        4    0.158

17    0.170       19    0.145               17    0.190       19    0.158

124    0.170       21    0.144               35    0.182       21    0.158

113    0.166       45    0.144               20    0.180       32    0.158

145    0.164        5    0.144              109    0.180      107    0.158

35    0.164        7    0.144              145    0.178      113    0.158

163    0.162      158    0.143               23    0.176      158    0.157

23    0.162       13    0.143               39    0.176        8    0.157

39    0.161      107    0.142              163    0.175       22    0.157

16    0.160        4    0.141                2    0.175       34    0.156

20    0.160      114    0.141               40    0.174       10    0.155

165    0.159      152    0.140              115    0.173      165    0.154

2    0.158        8    0.139                1    0.172      112    0.152

40    0.158      112    0.138               45    0.172      108    0.150

1    0.156       30    0.137              116    0.170       24    0.149

15    0.155       36    0.131               12    0.170      151    0.148

116    0.155      108    0.130               33    0.168      118    0.143

33    0.153      118    0.129              125    0.168       36    0.141

140    0.152      110    0.120              140    0.167       43    0.137

32    0.152       43    0.118               14    0.167      126    0.137

125    0.151      151    0.116               15    0.165      110    0.120

24    0.151      104    0.100               16    0.164      104    0.120

14    0.151      120    0.080               38    0.163      157    0.100

115    0.151      155    0.080               26    0.162      155    0.082

12    0.150      126    0.076              167    0.162      120    0.070 (*)

26    0.150                                152    0.160

Ammonium in precipitation                   Ammonium in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.267             Theoretical value:        0.321

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      67             Number of laboratories:      67

Arithmetic mean value:    0.272             Arithmetic mean value:    0.330

Median:                   0.268             Median:                   0.323

Standard deviation        0.044             Standard deviation        0.047

Rel. st. deviation (%)   16.037             Rel. st. deviation (%)   14.352

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      63

Arithmetic mean value:    0.276             Arithmetic mean value:    0.329

Median:                   0.269             Median:                   0.323

Standard deviation        0.030             Standard deviation        0.034

Rel. st. deviation (%)   10.751             Rel. st. deviation (%)   10.215

Results in decreasing order:                Results in decreasing order:

150    0.449 (*)    5    0.267              150    0.525 (*)  107    0.323

42    0.356        8    0.267               42    0.440 (*)    5    0.322

145    0.342       26    0.267              121    0.421       38    0.322

121    0.339       27    0.267              113    0.421       45    0.322

164    0.330       33    0.266              124    0.400        4    0.322

41    0.328      107    0.266               41    0.385      112    0.321

3    0.326      158    0.265              164    0.380      158    0.321

124    0.320      165    0.265                3    0.374      167    0.320

133    0.319      112    0.263               34    0.370      153    0.320

31    0.306       16    0.262              125    0.370      115    0.318

35    0.305       24    0.261               31    0.367       40    0.318

163    0.302      152    0.260               35    0.365       30    0.317

17    0.300       15    0.260              109    0.360       32    0.316

125    0.295       32    0.260               17    0.360       13    0.314

39    0.293      166    0.260               39    0.351       14    0.314

23    0.292      153    0.260               23    0.351        8    0.313

34    0.291      167    0.259              116    0.345      114    0.312

2    0.288       30    0.258                2    0.341       15    0.308

116    0.286      115    0.256               20    0.340       22    0.303

40    0.282       22    0.253              140    0.339      118    0.302

140    0.280      108    0.250               24    0.338       36    0.302

20    0.280      118    0.249              163    0.331      151    0.301

12    0.280      151    0.248               12    0.330      104    0.300

113    0.277       36    0.248              166    0.330      108    0.300

1    0.276       19    0.239                7    0.330      145    0.300

7    0.272       45    0.238              152    0.330      133    0.295

38    0.271      104    0.230               33    0.330       19    0.293

21    0.271      110    0.230               16    0.329      110    0.280

114    0.270       43    0.187               26    0.328      155    0.261

109    0.270      155    0.181 (*)           10    0.327      120    0.260

13    0.269      157    0.170 (*)            1    0.326      126    0.239

10    0.269      120    0.170 (*)          165    0.324       43    0.226 (*)

4    0.268      126    0.168 (*)           21    0.323      157    0.200 (*)

14    0.268                                 27    0.323

Table 21: Analytical results for pH in precipitations samples.

pH in precipitation                         pH in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        4.180             Theoretical value:        4.300

Unit: pH-unit                               Unit: pH-unit

Run 1:                                      Run 1:

Number of laboratories:      61             Number of laboratories:      61

Arithmetic mean value:    4.215             Arithmetic mean value:    4.310

Median:                   4.240             Median:                   4.340

Standard deviation        0.243             Standard deviation        0.244

Rel. st. deviation (%)    5.764             Rel. st. deviation (%)    5.659

Run 2:                                      Run 2:

Number of laboratories:      59             Number of laboratories:      60

Arithmetic mean value:    4.234             Arithmetic mean value:    4.340

Median:                   4.240             Median:                   4.340

Standard deviation        0.070             Standard deviation        0.068

Rel. st. deviation (%)    1.649             Rel. st. deviation (%)    1.575

Results in decreasing order:                Results in decreasing order:

104    4.790 (*)  124    4.240              133    4.450      158    4.340

133    4.370       33    4.240              109    4.440        8    4.340

110    4.360      153    4.230              110    4.440        4    4.340

109    4.350      121    4.230              115    4.440       36    4.340

34    4.323       20    4.220               19    4.430      164    4.330

19    4.320        4    4.220              104    4.430      121    4.330

115    4.310        8    4.220               30    4.430       21    4.330

30    4.300      158    4.220              108    4.410        1    4.320

10    4.300      114    4.220              112    4.410       32    4.320

23    4.300      107    4.220               14    4.410       20    4.320

112    4.300       36    4.220               34    4.391       16    4.320

108    4.290       21    4.220              124    4.380       22    4.320

14    4.280        1    4.210               10    4.380       12    4.320

151    4.280       16    4.210               23    4.380       13    4.310

140    4.280        7    4.210               31    4.380      116    4.310

165    4.270       22    4.210              140    4.370       45    4.310

39    4.270       32    4.210              113    4.370      114    4.310

116    4.260       12    4.200               39    4.370      165    4.300

113    4.260      155    4.200              153    4.370        5    4.300

35    4.250        5    4.190              163    4.370      155    4.300

126    4.250      120    4.190               26    4.360      150    4.300

150    4.250      166    4.190               27    4.360      120    4.300

46    4.250       45    4.185               46    4.360      166    4.300

13    4.250      118    4.180              126    4.360      118    4.290

26    4.250        3    4.167               33    4.360        3    4.275

31    4.250       17    4.100              152    4.350       17    4.240

152    4.240       38    4.060               35    4.350       38    4.180

163    4.240       40    4.020              151    4.348        2    4.110

164    4.240        2    3.980                7    4.340       40    4.080

15    4.240       24    2.510 (*)          107    4.340       24    2.510 (*)

27    4.240                                 15    4.340

pH in precipitation                         pH in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        4.520             Theoretical value:        4.480

Unit: pH-unit                               Unit: pH-unit

Run 1:                                      Run 1:

Number of laboratories:      61             Number of laboratories:      61

Arithmetic mean value:    4.527             Arithmetic mean value:    4.501

Median:                   4.560             Median:                   4.520

Standard deviation        0.278             Standard deviation        0.314

Rel. st. deviation (%)    6.139             Rel. st. deviation (%)    6.984

Run 2:                                      Run 2:

Number of laboratories:      60             Number of laboratories:      59

Arithmetic mean value:    4.561             Arithmetic mean value:    4.516

Median:                   4.560             Median:                   4.520

Standard deviation        0.084             Standard deviation        0.092

Rel. st. deviation (%)    1.838             Rel. st. deviation (%)    2.032

Results in decreasing order:                Results in decreasing order:

150    4.780        8    4.560              133    5.660 (*)   36    4.510

110    4.760       31    4.560              109    4.800        4    4.510

133    4.710      121    4.560              150    4.690       30    4.510

19    4.670       46    4.560               19    4.630      121    4.510

14    4.650       21    4.550              112    4.630      164    4.510

115    4.640      155    4.550              110    4.620      165    4.510

112    4.640       13    4.550               14    4.620       15    4.510

109    4.630       27    4.550              115    4.600        1    4.500

113    4.620      104    4.540               26    4.590       32    4.500

34    4.611       32    4.540              108    4.590       21    4.500

26    4.610       36    4.540              140    4.580      114    4.500

140    4.610      107    4.540              151    4.575       16    4.490

108    4.610      114    4.530              113    4.570       45    4.490

30    4.610       16    4.530               39    4.570      155    4.490

124    4.600       20    4.530               22    4.560      107    4.490

33    4.590       45    4.525               35    4.560       12    4.490

152    4.590       12    4.520               10    4.550       20    4.480

163    4.590      120    4.520               34    4.549       13    4.480

164    4.590      165    4.520              163    4.540      120    4.480

39    4.590      166    4.520              126    4.540      104    4.470

126    4.590        5    4.510               33    4.540      166    4.470

35    4.580      118    4.510              152    4.530      118    4.470

23    4.580        3    4.491              153    4.530        3    4.467

22    4.580       17    4.490               27    4.530        5    4.460

7    4.580      151    4.465               46    4.530       17    4.380

10    4.580      116    4.450              124    4.520      116    4.380

15    4.570       38    4.410              158    4.520       38    4.320

1    4.570       40    4.300               31    4.520       40    4.220

153    4.560        2    4.240                8    4.520        2    4.190

158    4.560       24    2.490 (*)            7    4.520       24    2.480 (*)

4    4.560                                 23    4.520

Table 22: Analytical results for strong acid calculated from pH.

Strong acid  calculated from pH             Strong acid  calculated from pH

Sample no.: G1                              Sample no.: G2

Theoretical value:       67.000             Theoretical value:       50.000

Unit: µeq/l                                 Unit: µeq/l

Run 1:                                      Run 1:

Number of laboratories:      58             Number of laboratories:      58

Arithmetic mean value:  110.776             Arithmetic mean value:   98.928

Median:                  57.540             Median:                  45.710

Standard deviation      398.282             Standard deviation      399.767

Rel. st. deviation (%)  359.539             Rel. st. deviation (%)  404.098

Run 2:                                      Run 2:

Number of laboratories:      57             Number of laboratories:      57

Arithmetic mean value:   58.503             Arithmetic mean value:   46.448

Median:                  57.540             Median:                  45.710

Standard deviation       12.323             Standard deviation        8.547

Rel. st. deviation (%)   21.064             Rel. st. deviation (%)   18.402

Results in decreasing order:                Results in decreasing order:

24 3090.300 (*)  163   57.540               24 3090.300 (*)  158   45.710

2  104.710       27   57.540               40   83.180        4   45.710

40   95.500      152   57.540                2   77.620        7   45.710

38   87.100       33   57.540               38   66.070      151   44.870

17   79.430      150   56.230               17   57.540      152   44.670

3   68.080       13   56.230                3   53.090       35   44.670

118   66.070       35   56.230              118   51.290      126   43.650

45   65.310       31   56.230              120   50.120       26   43.650

120   64.570      126   56.230                5   50.120       33   43.650

5   64.570       26   56.230              155   50.120       27   43.650

166   64.570      116   54.950              166   50.120      140   42.660

155   63.100      113   54.950              165   50.120      153   42.660

12   63.100      165   53.700              150   50.120       39   42.660

1   61.660       39   53.700               13   48.980      163   42.660

7   61.660       14   52.480               45   48.980      113   42.660

32   61.660      140   52.480              114   48.980       31   41.690

22   61.660      151   52.480              116   48.980       10   41.690

16   61.660      108   51.290               16   47.860       23   41.690

4   60.260       23   50.120               22   47.860       34   40.640

8   60.260       30   50.120                1   47.860       14   38.900

21   60.260      112   50.120               32   47.860      108   38.900

158   60.260       10   50.120               20   47.860      112   38.900

20   60.260      115   48.980               12   47.860      104   37.150

36   60.260       19   47.860              121   46.770       19   37.150

114   60.260       34   47.530              164   46.770       30   37.150

121   58.880      109   44.670               21   46.770      115   36.310

153   58.880      110   43.650                8   45.710      110   36.310

164   57.540      133   42.660               15   45.710      109   36.310

15   57.540      104   16.220               36   45.710      133   35.480

Strong acid  calculated from pH             Strong acid  calculated from pH

Sample no.: G3                              Sample no.: G4

Theoretical value:       30.000             Theoretical value:       33.000

Unit: µeq/l                                 Unit: µeq/l

Run 1:                                      Run 1:

Number of laboratories:      58             Number of laboratories:      58

Arithmetic mean value:   83.364             Arithmetic mean value:   87.302

Median:                  27.540             Median:                  30.550

Standard deviation      421.263             Standard deviation      430.845

Rel. st. deviation (%)  505.332             Rel. st. deviation (%)  493.512

Run 2:                                      Run 2:

Number of laboratories:      57             Number of laboratories:      57

Arithmetic mean value:   28.055             Arithmetic mean value:   30.740

Median:                  27.540             Median:                  30.200

Standard deviation        6.340             Standard deviation        8.635

Rel. st. deviation (%)   22.597             Rel. st. deviation (%)   28.089

Results in decreasing order:                Results in decreasing order:

24 3235.940 (*)   31   27.540               24 3311.310 (*)    7   30.200

2   57.540        4   27.540                2   64.570        8   30.200

40   50.120       15   26.920               40   60.260       23   30.200

38   38.900        1   26.920               38   47.860       31   30.200

116   35.480       35   26.300              116   41.690      158   30.200

151   34.280       10   26.300               17   41.690      152   29.510

17   32.360       23   26.300                5   34.670       27   29.510

3   32.280       22   26.300                3   34.120      153   29.510

118   30.900        7   26.300              104   33.880      163   28.840

5   30.900       39   25.700              118   33.880       33   28.840

12   30.200      163   25.700              166   33.880      126   28.840

166   30.200      164   25.700              120   33.110       34   28.250

165   30.200      126   25.700               20   33.110       10   28.180

120   30.200      152   25.700               13   33.110       35   27.540

45   29.850       33   25.700               16   32.360       22   27.540

114   29.510       30   24.550               45   32.360      113   26.920

20   29.510      140   24.550               12   32.360       39   26.920

16   29.510      108   24.550              155   32.360      151   26.610

32   28.840       26   24.550                1   31.620      140   26.300

36   28.840       34   24.490               21   31.620       26   25.700

104   28.840      113   23.990              114   31.620      108   25.700

155   28.180      109   23.440               32   31.620      115   25.120

13   28.180      115   22.910                4   30.900      110   23.990

27   28.180      112   22.910              121   30.900       14   23.990

21   28.180       14   22.390              164   30.900      112   23.440

153   27.540       19   21.380              165   30.900       19   23.440

8   27.540      133   19.500               15   30.900      150   20.420

158   27.540      110   17.380               30   30.900      109   15.850

121   27.540      150   16.600               36   30.900      133    2.190

Table 23: Analytical results for chloride in precipitations samples.

Chloride in precipitation                   Chloride in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.154             Theoretical value:        0.174

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    0.210             Arithmetic mean value:    0.229

Median:                   0.156             Median:                   0.170

Standard deviation        0.369             Standard deviation        0.367

Rel. st. deviation (%)  175.577             Rel. st. deviation (%)  160.303

Run 2:                                      Run 2:

Number of laboratories:      61             Number of laboratories:      61

Arithmetic mean value:    0.164             Arithmetic mean value:    0.183

Median:                   0.156             Median:                   0.170

Standard deviation        0.059             Standard deviation        0.061

Rel. st. deviation (%)   36.158             Rel. st. deviation (%)   33.523

Results in decreasing order:                Results in decreasing order:

40    3.033 (*)  158    0.156               40    3.033 (*)   30    0.170

17    0.466      125    0.153               43    0.468      150    0.170

108    0.330       31    0.152               17    0.343      166    0.170

19 <  0.31

110 <  0.30

133    0.277      145    0.152              108    0.320       31    0.170

109 <  0.25

42    0.238        3    0.152              133    0.313       20    0.170

19 <  0.31

110 <  0.30

120    0.230        1    0.152               42    0.282       38    0.170

109 <  0.25

115    0.218      140    0.150                2    0.238      124    0.170

2    0.201       20    0.150              145    0.237        3    0.169

152    0.200      166    0.150              165    0.220       26    0.167

165    0.199      124    0.150              104    0.220       23    0.167

22    0.196       16    0.148              115    0.208       35    0.166

36    0.190       23    0.147               39    0.201       15    0.165

104    0.180       21    0.142              152    0.200       45    0.163

38    0.180      163    0.142               22    0.197       16    0.162

39    0.178       12    0.140               36    0.196       21    0.161

167    0.176      114    0.140                7    0.191       12    0.160

7    0.175       35    0.140              167    0.191      140    0.160

13    0.174       45    0.140              120    0.190      163    0.160

113    0.172       10    0.137              114    0.190      116    0.158

30    0.170      116    0.137               13    0.189       10    0.155

150    0.170       43    0.137              151    0.185      153    0.150

8    0.165      107    0.131               33    0.180      107    0.146

121    0.164      153    0.130                8    0.179       41    0.142

32    0.163       41    0.127               32    0.179      118    0.140

151    0.161      126    0.120              125    0.177      126    0.140

33    0.160      164    0.120                5    0.177      113    0.137

5    0.159       14    0.112              121    0.175      164    0.130

26    0.157      118    0.110               27    0.174       14    0.128

15    0.157      155    0.100              112    0.173      155    0.101

112    0.156       24    0.038              158    0.173       34    0.067

27    0.156       34    0.027                1    0.170       24    0.043

Chloride in precipitation                   Chloride in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.290             Theoretical value:        0.309

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    0.331             Arithmetic mean value:    0.355

Median:                   0.287             Median:                   0.300

Standard deviation        0.361             Standard deviation        0.366

Rel. st. deviation (%)  108.950             Rel. st. deviation (%)  103.191

Run 2:                                      Run 2:

Number of laboratories:      63             Number of laboratories:      63

Arithmetic mean value:    0.287             Arithmetic mean value:    0.310

Median:                   0.286             Median:                   0.300

Standard deviation        0.063             Standard deviation        0.082

Rel. st. deviation (%)   21.784             Rel. st. deviation (%)   26.450

Results in decreasing order:                Results in decreasing order:

40    3.133 (*)    1    0.286               40    3.167 (*)   33    0.300

108    0.560        3    0.285               43    0.636      150    0.300

110 <  0.30

43    0.400       32    0.285               17    0.620        3    0.299

17    0.373       23    0.283              108    0.540       21    0.299

133    0.369       31    0.280               42    0.442      125    0.298

2    0.361      166    0.280              133    0.406       31    0.298

165    0.355       33    0.280                2    0.367      140    0.290

104    0.350       45    0.278              112    0.361      114    0.290

112    0.345       30    0.277              104    0.360      166    0.290

42    0.338       21    0.275               13    0.359       12    0.290

113    0.330       16    0.275              165    0.352      113    0.289

13    0.321      163    0.274               39    0.346      163    0.287

151    0.316      140    0.270              151    0.337       35    0.284

36    0.315       12    0.270              145    0.331        4    0.282

39    0.313      114    0.270               36    0.330      118    0.280

4    0.311       35    0.268              167    0.320       15    0.280

19    0.310       15    0.268              115    0.320       45    0.279

115    0.303      121    0.261               38    0.320       14    0.275

167    0.302      118    0.260               32    0.316      121    0.275

125    0.302      153    0.260               27    0.314      116    0.272

124    0.300       10    0.258               30    0.312       10    0.271

22    0.300      116    0.252                5    0.311      126    0.270

110 <  0.30

7    0.294      120    0.250               26    0.311      153    0.270

109 <  0.25

27    0.294      107    0.241               20    0.310      109    0.260

5    0.293      164    0.240              124    0.310      107    0.254

19 <  0.31

26    0.293      126    0.240                7    0.309      164    0.250

38    0.290       41    0.232              158    0.307       41    0.242

145    0.290       14    0.220               23    0.305      120    0.240

150    0.290      152    0.200                1    0.304      152    0.200

158    0.290      155    0.184                8    0.304      155    0.197

20    0.290       24    0.098               22    0.302       34    0.152

8    0.288       34    0.094               16    0.300       24    0.123

Table 24:	Analytical results for sodium in precipitations samples.

Sodium in precipitation                     Sodium in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.243             Theoretical value:        0.274

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.238             Arithmetic mean value:    0.261

Median:                   0.246             Median:                   0.274

Standard deviation        0.040             Standard deviation        0.050

Rel. st. deviation (%)   17.026             Rel. st. deviation (%)   19.185

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      62

Arithmetic mean value:    0.243             Arithmetic mean value:    0.270

Median:                   0.247             Median:                   0.275

Standard deviation        0.024             Standard deviation        0.029

Rel. st. deviation (%)   10.002             Rel. st. deviation (%)   10.712

Results in decreasing order:                Results in decreasing order:

40    0.308      125    0.246               40    0.355        8    0.272

20    0.290       26    0.245               41    0.316       32    0.272

165    0.289      113    0.244               20    0.300      114    0.270

124    0.270        8    0.244              151    0.300      166    0.270

110    0.267       36    0.243              165    0.295       16    0.270

39    0.266       14    0.243               39    0.294       23    0.269

145    0.266       23    0.241               24    0.292       36    0.268

4    0.263      114    0.240               45    0.292        1    0.266

126    0.260        1    0.239               19    0.290       15    0.266

108    0.260        3    0.238              126    0.290      152    0.265

150    0.258      152    0.238              108    0.290        3    0.264

41    0.258       32    0.238                7    0.290      107    0.262

7    0.258       22    0.236              124    0.290      113    0.260

19    0.258       15    0.236               21    0.288       12    0.260

112    0.257      107    0.235                4    0.287        2    0.258

5    0.256        2    0.233              145    0.287       33    0.258

30    0.256       33    0.231              112    0.285       31    0.255

116    0.255       12    0.230              163    0.284      164    0.250

163    0.254      157    0.230              110    0.283      115    0.244

158    0.253       31    0.228              150    0.283       22    0.241

155    0.252      115    0.228                5    0.282       10    0.241

24    0.251      104    0.223               14    0.282      104    0.233

45    0.250       10    0.217               30    0.282       35    0.231

153    0.250      118    0.210              140    0.280      157    0.230

13    0.250      164    0.210              120    0.280      167    0.226

16    0.250       35    0.209              153    0.280       42    0.224

140    0.250       42    0.203              116    0.280      118    0.220

120    0.250      167    0.195              158    0.280       38    0.180

166    0.250      109    0.170               13    0.279      109    0.170

21    0.249       38    0.160              155    0.278      133    0.150 (*)

151    0.248      133    0.100 (*)           27    0.275       17    0.053 (*)

17 <  0.025

27    0.247      121    0.023 (*)          125    0.274      121    0.039 (*)

26    0.274

Sodium in precipitation                     Sodium in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.385             Theoretical value:        0.433

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.368             Arithmetic mean value:    0.432

Median:                   0.380             Median:                   0.432

Standard deviation        0.052             Standard deviation        0.104

Rel. st. deviation (%)   14.029             Rel. st. deviation (%)   24.110

Run 2:                                      Run 2:

Number of laboratories:      61             Number of laboratories:      61

Arithmetic mean value:    0.378             Arithmetic mean value:    0.425

Median:                   0.382             Median:                   0.432

Standard deviation        0.028             Standard deviation        0.033

Rel. st. deviation (%)    7.338             Rel. st. deviation (%)    7.862

Results in decreasing order:                Results in decreasing order:

165    0.450      140    0.380              157    1.110 (*)  152    0.431

40    0.444      114    0.380               40    0.659 (*)  114    0.430

20    0.420      108    0.380              165    0.471       21    0.430

150    0.420       36    0.379               20    0.470      116    0.430

39    0.416       32    0.379               39    0.470      140    0.430

7    0.410      125    0.378               41    0.463      125    0.428

41    0.401        1    0.375                7    0.461       36    0.425

4    0.401       16    0.371              124    0.460        1    0.423

166    0.400       12    0.370              110    0.455       15    0.421

124    0.400        2    0.369                4    0.453      151    0.421

30    0.400      107    0.369               30    0.451       35    0.421

112    0.398       15    0.367              163    0.450       12    0.420

5    0.395        3    0.364              126    0.450      108    0.420

110    0.395      151    0.364              155    0.449       16    0.420

155    0.394      145    0.360              112    0.446      109    0.420

13    0.393       24    0.360                5    0.445      107    0.414

158    0.393       31    0.357              113    0.445       24    0.410

163    0.391      115    0.353              150    0.444        3    0.409

120    0.390      157    0.350               23    0.443       33    0.404

126    0.390      167    0.349              158    0.442       31    0.403

19    0.388       35    0.345               32    0.441      167    0.396

27    0.387       10    0.343               19    0.441      115    0.393

23    0.386       22    0.342              166    0.440       10    0.391

45    0.385      164    0.340              153    0.440      164    0.390

113    0.384      104    0.333              120    0.440      104    0.378

21    0.383      118    0.320               27    0.438       42    0.378

26    0.383       42    0.308                2    0.438      118    0.370

8    0.383      133    0.300               45    0.438       22    0.366

116    0.383       38    0.254 (*)           14    0.438       38    0.300

14    0.383      109    0.250 (*)          145    0.437      133    0.300

33    0.382      121    0.167 (*)           13    0.435      121    0.208 (*)

152    0.382       17    0.144 (*)            8    0.435       17    0.180 (*)

153    0.380                                 26    0.432

Table 25:	Analytical results for magnesium in precipitations samples.

Magnesium in precipitation                  Magnesium in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.175             Theoretical value:        0.072

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.177             Arithmetic mean value:    0.074

Median:                   0.172             Median:                   0.070

Standard deviation        0.029             Standard deviation        0.027

Rel. st. deviation (%)   16.432             Rel. st. deviation (%)   36.541

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      64

Arithmetic mean value:    0.174             Arithmetic mean value:    0.071

Median:                   0.172             Median:                   0.070

Standard deviation        0.017             Standard deviation        0.013

Rel. st. deviation (%)    9.555             Rel. st. deviation (%)   17.772

Results in decreasing order:                Results in decreasing order:

145    0.304 (*)  104    0.172               40    0.266 (*)   19    0.070

157    0.290 (*)   26    0.171              145    0.123      109    0.070

20    0.230       21    0.171               20    0.100      126    0.070

15    0.207      116    0.171               15    0.094      151    0.069

165    0.203       30    0.171                2    0.093      121    0.069

153    0.200       13    0.171               23    0.090      133    0.069

35    0.200      120    0.170              107    0.089      116    0.068

114    0.200      109    0.170               35    0.085        7    0.068

163    0.199      140    0.170               24    0.084       30    0.068

2    0.193      166    0.169              163    0.084       36    0.068

40    0.193      151    0.169              165    0.082       14    0.068

24    0.190       16    0.168              125    0.081       21    0.067

107    0.188      150    0.168              153    0.080      166    0.067

23    0.188      112    0.167              114    0.080      113    0.066

10    0.184      118    0.166              157    0.080      118    0.066

42    0.184      113    0.165              124    0.080       34    0.066

8    0.182       36    0.165               10    0.079      112    0.065

125    0.180       33    0.165               22    0.075      150    0.064

155    0.179      167    0.163                8    0.075      152    0.064

27    0.179      152    0.160               16    0.075       12    0.064

12    0.179      124    0.160              155    0.074       38    0.062

5    0.179       34    0.158               31    0.074      167    0.062

22    0.178       14    0.157                3    0.074      115    0.062

7    0.177        1    0.156               27    0.073        1    0.061

110    0.177        4    0.154               45    0.073       41    0.058

3    0.176      115    0.153               42    0.072        4    0.056

133    0.176      158    0.151               39    0.072      104    0.055

19    0.176      108    0.150                5    0.072      110    0.052

31    0.176       38    0.149               13    0.071      158    0.051

121    0.174       17    0.145              120    0.071      108    0.050

41    0.174      126    0.130               26    0.071      164    0.050

39    0.174      164    0.090 (*)           33    0.071       17    0.044

45    0.172                                140    0.070

Magnesium in precipitation                  Magnesium in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.088             Theoretical value:        0.114

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.097             Arithmetic mean value:    0.119

Median:                   0.088             Median:                   0.110

Standard deviation        0.077             Standard deviation        0.050

Rel. st. deviation (%)   80.109             Rel. st. deviation (%)   41.851

Run 2:                                      Run 2:

Number of laboratories:      64             Number of laboratories:      63

Arithmetic mean value:    0.087             Arithmetic mean value:    0.110

Median:                   0.087             Median:                   0.110

Standard deviation        0.012             Standard deviation        0.012

Rel. st. deviation (%)   14.001             Rel. st. deviation (%)   10.507

Results in decreasing order:                Results in decreasing order:

40    0.704 (*)   31    0.087              157    0.400 (*)  121    0.110

107    0.112      104    0.087               40    0.377 (*)  126    0.110

15    0.110       13    0.086              107    0.135      109    0.110

2    0.110       42    0.086               35    0.134      166    0.110

20    0.110       21    0.085                2    0.132       26    0.110

114    0.110      151    0.085              114    0.130       31    0.110

35    0.105      116    0.085               20    0.130      133    0.109

165    0.103       36    0.085              163    0.127      116    0.109

125    0.103      121    0.085              165    0.127      151    0.108

163    0.102      120    0.085               15    0.125       21    0.108

23    0.102       34    0.085               42    0.124      120    0.108

145    0.100      118    0.084              125    0.123       36    0.108

10    0.094      166    0.084               23    0.122      104    0.107

24    0.094       12    0.084               24    0.121      145    0.106

22    0.092      133    0.084              124    0.120      118    0.105

8    0.092       14    0.083               10    0.118       13    0.105

27    0.091      167    0.083                8    0.118      112    0.105

155    0.091      112    0.083               16    0.116      152    0.104

7    0.091      150    0.081              155    0.115       41    0.103

5    0.091      152    0.080               19    0.114      167    0.102

124    0.090      113    0.080               27    0.114       14    0.102

140    0.090        1    0.078                5    0.113      150    0.101

19    0.090       41    0.078               39    0.112      113    0.100

153    0.090      115    0.077                3    0.112        1    0.099

157    0.090       38    0.077               45    0.112       38    0.097

109    0.090        4    0.072               22    0.112      115    0.097

3    0.089      110    0.071               33    0.112      110    0.095

39    0.089      158    0.069               12    0.111        4    0.094

45    0.088       17    0.061               34    0.111      158    0.091

16    0.088      108    0.060               30    0.111      108    0.090

33    0.088      164    0.060                7    0.110      164    0.080

26    0.088      126    0.050              153    0.110       17    0.079

30    0.088                                140    0.110

Table 26:	Analytical results for calcium in precipitations samples.

Calcium in precipitation                    Calcium in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.128             Theoretical value:        0.166

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.155             Arithmetic mean value:    0.167

Median:                   0.130             Median:                   0.170

Standard deviation        0.113             Standard deviation        0.038

Rel. st. deviation (%)   72.731             Rel. st. deviation (%)   22.846

Run 2:                                      Run 2:

Number of laboratories:      62             Number of laboratories:      59

Arithmetic mean value:    0.137             Arithmetic mean value:    0.171

Median:                   0.130             Median:                   0.170

Standard deviation        0.035             Standard deviation        0.025

Rel. st. deviation (%)   25.429             Rel. st. deviation (%)   14.565

Results in decreasing order:                Results in decreasing order:

157    0.930 (*)   30    0.130              107    0.286 (*)   39    0.169

20    0.480 (*)    5    0.130               35    0.220      112    0.167

126    0.240       14    0.130              124    0.220       30    0.167

22    0.228       39    0.130               20    0.220       42    0.166

124    0.210      112    0.130              163    0.218       21    0.165

145 <  0.2

165    0.192       21    0.129              153    0.210       36    0.165

110    0.187       45    0.129              126    0.210        4    0.165

10    0.185      125    0.128               41    0.202       33    0.164

145  < 0.2

40    0.185       15    0.128               10    0.199       45    0.163

109    0.180       24    0.128               23    0.196      116    0.161

140    0.180       16    0.128              113    0.194      118    0.160

153    0.170      120    0.128               22    0.190       14    0.160

7    0.168       31    0.127              109    0.190       13    0.160

163    0.167      116    0.126              140    0.190      166    0.160

23    0.160        3    0.125               34    0.188      115    0.157

8    0.154       36    0.125              165    0.187      125    0.156

41    0.150        4    0.125               24    0.186        3    0.156

113    0.150        2    0.124              151    0.185      110    0.152

152    0.150       33    0.124                2    0.184       31    0.150

34    0.145       13    0.123                7    0.182       17    0.144

121    0.144      104    0.121                8    0.182       15    0.141

133    0.143      166    0.120                5    0.181      167    0.140

38    0.142      115    0.120               19    0.180       12    0.140

35    0.142      167    0.108               38    0.179      114    0.130

42    0.138       12    0.100              155    0.177        1    0.129

151    0.137      114    0.100               26    0.176      157    0.120

107    0.137        1    0.092              121    0.172       32    0.120

155    0.134       17    0.086               27    0.171      150    0.112

26    0.133      150    0.083               16    0.170      158    0.090 (*)

19    0.133       32    0.076              120    0.170       40    0.072 (*)

27    0.132      158    0.057              104    0.170      164    0.070 (*)

118    0.130      164    0.050              152    0.170      133    0.042 (*)

Calcium in precipitation                    Calcium in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.192             Theoretical value:        0.230

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      66             Number of laboratories:      66

Arithmetic mean value:    0.192             Arithmetic mean value:    0.236

Median:                   0.195             Median:                   0.233

Standard deviation        0.038             Standard deviation        0.076

Rel. st. deviation (%)   19.542             Rel. st. deviation (%)   32.103

Run 2:                                      Run 2:

Number of laboratories:      60             Number of laboratories:      63

Arithmetic mean value:    0.194             Arithmetic mean value:    0.231

Median:                   0.196             Median:                   0.233

Standard deviation        0.024             Standard deviation        0.044

Rel. st. deviation (%)   12.150             Rel. st. deviation (%)   18.969

Results in decreasing order:                Results in decreasing order:

107    0.324 (*)   39    0.194              157    0.710 (*)  112    0.233

35    0.273 (*)   16    0.193               34    0.386       27    0.233

108    0.250       45    0.191              107    0.360       13    0.231

163    0.244      112    0.191               35    0.309       21    0.231

167    0.232       36    0.191              163    0.285      166    0.230

109    0.230        4    0.190              113    0.285      140    0.230

124    0.230       13    0.186              124    0.280        2    0.229

34    0.228       42    0.185              108    0.280       36    0.226

10    0.222      120    0.185               10    0.264        4    0.225

126    0.220       21    0.185              126    0.260       14    0.220

20    0.220       15    0.183              151    0.260       33    0.220

151    0.218      116    0.183               20    0.260       15    0.216

41    0.214      125    0.182               42    0.256      110    0.216

7    0.214      118    0.180               38    0.254      125    0.214

23    0.211      166    0.180                8    0.251        3    0.213

152    0.210       32    0.180              153    0.250      115    0.212

153    0.210      104    0.179               41    0.244      104    0.212

8    0.209       22    0.179                7    0.243      165    0.211

19    0.205      115    0.176               23    0.242      118    0.210

38    0.202      110    0.174              152    0.240      116    0.209

14    0.200        3    0.173               16    0.240       31    0.201

113    0.200      157    0.170              120    0.240       32    0.200

140    0.200       31    0.165              109    0.240       12    0.200

145 <  0.2                                                    145 <  0.2

24    0.199       17    0.164               30    0.238       17    0.199

30    0.199       12    0.160               45    0.237        1    0.190

165    0.198        1    0.156               24    0.236      158    0.172

121    0.198      114    0.150               26    0.236      114    0.170

2    0.197      150    0.134              121    0.236      150    0.169

5    0.197      158    0.134              167    0.236       22    0.157

155    0.196       40    0.105 (*)            5    0.235      164    0.110

26    0.196      164    0.090 (*)           19    0.234       40    0.106

33    0.196      133    0.068 (*)           39    0.234      133    0.083 (*)

27    0.195                                155    0.233

Table 27:	Analytical results for potassium in precipitations samples.

Potassium in precipitation                  Potassium in precipitation

Sample no.: G1                              Sample no.: G2

Theoretical value:        0.204             Theoretical value:        0.187

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      65             Number of laboratories:      65

Arithmetic mean value:    0.197             Arithmetic mean value:    0.180

Median:                   0.198             Median:                   0.180

Standard deviation        0.026             Standard deviation        0.024

Rel. st. deviation (%)   13.127             Rel. st. deviation (%)   13.223

Run 2:                                      Run 2:

Number of laboratories:      59             Number of laboratories:      56

Arithmetic mean value:    0.195             Arithmetic mean value:    0.182

Median:                   0.197             Median:                   0.181

Standard deviation        0.017             Standard deviation        0.017

Rel. st. deviation (%)    8.832             Rel. st. deviation (%)    9.400

Results in decreasing order:                Results in decreasing order:

121    0.314 (*)  158    0.197              121    0.237 (*)  112    0.180

140    0.270 (*)   24    0.196              140    0.230 (*)  158    0.180

150    0.240       23    0.195              150    0.220       16    0.180

10    0.234       42    0.193              153    0.220       12    0.180

2    0.233       13    0.193                2    0.219      124    0.180

20    0.220       45    0.191               10    0.215       15    0.179

5    0.217       31    0.191               41    0.202       33    0.178

22    0.215      126    0.190               40    0.201      107    0.178

41    0.211       12    0.190              125    0.200      152    0.178

19    0.210      107    0.190              120    0.200       24    0.178

110 <  0.200

108    0.210      157    0.190                5    0.199        8    0.176

116    0.209       16    0.190               19    0.194       36    0.175

7    0.208      153    0.190              116    0.193       30    0.173

113    0.205        1    0.189                3    0.193        1    0.173

151    0.205      145    0.189               31    0.192      114    0.170

125    0.204        8    0.189               32    0.191        4    0.170

39    0.204       33    0.188              133    0.190      126    0.170

36    0.203       14    0.187              108    0.190      104    0.169

155    0.202       30    0.187               20    0.190       42    0.166

163    0.202      104    0.186                7    0.188       22    0.163

3    0.202      120    0.180               45    0.188       14    0.162

15    0.201        4    0.180               39    0.187      165    0.157

112    0.201      167    0.177              151    0.187      167    0.150

21    0.201      164    0.170              163    0.187      157    0.150

26    0.200      118    0.170               26    0.184      118    0.150

133    0.200       38    0.164               21    0.184      164    0.150

32    0.200       17    0.163               23    0.184       38    0.144

124    0.200      115    0.160               27    0.183      145    0.132 (*)

110 <  0.2

27    0.199      114    0.160              155    0.182       17    0.127 (*)

152    0.198      165    0.152               13    0.182      115    0.124 (*)

40    0.198       35    0.130 (*)          113    0.181       35    0.115 (*)

109 <  0.08                                 109 <  0.08

Potassium in precipitation                  Potassium in precipitation

Sample no.: G3                              Sample no.: G4

Theoretical value:        0.153             Theoretical value:        0.119

Unit: µg/l                                  Unit: µg/l

Run 1:                                      Run 1:

Number of laboratories:      62             Number of laboratories:      64

Arithmetic mean value:    0.149             Arithmetic mean value:    0.160

Median:                   0.148             Median:                   0.117

Standard deviation        0.018             Standard deviation        0.311

Rel. st. deviation (%)   12.050             Rel. st. deviation (%)  194.317

Run 2:                                      Run 2:

Number of laboratories:      54             Number of laboratories:      57

Arithmetic mean value:    0.148             Arithmetic mean value:    0.119

Median:                   0.148             Median:                   0.116

Standard deviation        0.013             Standard deviation        0.020

Rel. st. deviation (%)    9.037             Rel. st. deviation (%)   16.798

Results in decreasing order:                Results in decreasing order:

166 <  0.41

150    0.210 (*)    8    0.148              157    2.480 (*)   45    0.116

166 <  0.41

110 <  0.200                                109    0.200 (*)

113    0.195 (*)  112    0.148               40    0.191       27    0.116

2    0.187 (*)   16    0.148              113    0.190       15    0.115

140    0.180      152    0.147              165    0.177       33    0.114

10    0.174       21    0.147              140    0.150      158    0.113

121  < 0.15

153    0.170      158    0.146              150    0.140        1    0.113

19    0.170       33    0.146               10    0.135       21    0.113

41    0.167       45    0.146               41    0.132       31    0.112

165    0.166       24    0.142                2    0.132      112    0.112

133    0.160       31    0.141               19    0.132      152    0.111

20    0.160       30    0.141               20    0.130       12    0.110

108    0.160      107    0.140              108    0.130        4    0.110

5    0.159       14    0.140              124    0.130      126    0.110

116    0.158      114    0.140               23    0.129       13    0.110

7    0.156       12    0.140              116    0.123      114    0.110

23    0.155        4    0.140               42    0.122      118    0.110

32    0.155      126    0.140               39    0.121       30    0.109

155    0.154        1    0.140                7    0.121      151    0.109

125    0.154      124    0.140                5    0.121      107    0.108

15    0.154       36    0.136              133    0.120       36    0.106

39    0.153       42    0.132              153    0.120      104    0.104

151    0.152       40    0.132                3    0.120      167    0.103

104    0.151      167    0.130              125    0.120       14    0.103

3    0.151      118    0.130              155    0.119       24    0.102

27    0.150       38    0.120               16    0.118      120    0.100

145 <  0.1

13    0.150       22    0.120               26    0.118       38    0.096

120    0.150      157    0.120              163    0.118      164    0.090

163    0.150      164    0.120                8    0.118       22    0.088

121 <  0.15

26    0.149      115    0.108 (*)           32    0.117      115    0.085

109 <  0.08                                 109 <  0.08

17 <  0.077

Table 28:	Analytical results for conductivity in precipitations samples.

Conductivity in precipitation               Conductivity in precipitation

Sample no.: 1                               Sample no.: 2

Theoretical value:       35.000             Theoretical value:       27.000

Unit: µS/cm                                 Unit: µS/cm

Run 1:                                      Run 1:

Number of laboratories:      61             Number of laboratories:      61

Arithmetic mean value:   33.100             Arithmetic mean value:   26.279

Median:                  33.500             Median:                  26.600

Standard deviation        2.453             Standard deviation        2.179

Rel. st. deviation (%)    7.411             Rel. st. deviation (%)    8.294

Run 2:                                      Run 2:

Number of laboratories:      59             Number of laboratories:      59

Arithmetic mean value:   33.376             Arithmetic mean value:   26.547

Median:                  33.500             Median:                  26.700

Standard deviation        1.885             Standard deviation        1.545

Rel. st. deviation (%)    5.649             Rel. st. deviation (%)    5.821

Results in decreasing order:                Results in decreasing order:

24   37.200      115   33.450              116   30.140      165   26.500

126   37.100      107   33.400              126   29.400      163   26.500

5   36.600        8   33.300               14   28.800      155   26.480

14   36.400      164   33.110                5   28.700        8   26.400

166   36.000      114   32.900               24   28.600       16   26.300

21   35.440       30   32.800                4   28.300      109   26.300

38   35.400       17   32.750              166   28.300      115   26.200

12   35.400       45   32.710              121   28.200      164   26.110

39   35.400        1   32.700               21   28.020      152   26.100

20   35.300      165   32.400               12   28.000      151   26.100

19   35.150      153   32.360                3   28.000        1   26.100

3   35.000      151   32.100              120   28.000      118   26.000

120   35.000       23   32.000               19   27.900      153   25.710

46   34.970      118   32.000               38   27.800      110   25.500

10   34.900      152   32.000               15   27.700       45   25.457

31   34.700      150   31.900               31   27.600      124   25.400

15   34.700      110   31.700               30   27.600       34   25.200

121   34.400      112   31.500               10   27.400      113   25.100

33   34.200       34   31.400               33   27.300      112   25.000

158   34.200       36   31.300               17   27.250      150   25.000

4   34.100      108   31.200                2   27.000       36   24.400

2   34.000      116   31.100              158   27.000      108   24.400

13   34.000       40   31.000               13   27.000      140   24.300

109   33.900      113   30.600              114   27.000       40   24.000

124   33.800      145   30.100               46   26.940      104   24.000

155   33.800      140   29.900               27   26.800      133   23.590

35   33.700        7   29.560               35   26.720        7   23.460

163   33.700       22   28.800               32   26.700      145   22.500

27   33.700      133   27.890 (*)           23   26.700       22   21.300 (*)

32   33.500      104   22.000 (*)          107   26.700       20   15.440 (*)

16   33.500                                 39   26.600

Conductivity in precipitation               Conductivity in precipitation

Sample no.: 3                               Sample no.: 4

Theoretical value:       21.000             Theoretical value:       23.000

Unit: µS/cm                                 Unit: µS/cm

Run 1:                                      Run 1:

Number of laboratories:      60             Number of laboratories:      60

Arithmetic mean value:   20.263             Arithmetic mean value:   22.921

Median:                  20.450             Median:                  23.050

Standard deviation        1.222             Standard deviation        1.927

Rel. st. deviation (%)    6.028             Rel. st. deviation (%)    8.406

Run 2:                                      Run 2:

Number of laboratories:      56             Number of laboratories:      57

Arithmetic mean value:   20.370             Arithmetic mean value:   23.025

Median:                  20.500             Median:                  23.100

Standard deviation        0.978             Standard deviation        1.182

Rel. st. deviation (%)    4.803             Rel. st. deviation (%)    5.135

Results in decreasing order:                Results in decreasing order:

45   23.179 (*)  158   20.400               13   30.000 (*)  163   23.000

126   22.700      153   20.380              126   26.100       32   23.000

13   22.000      116   20.260              116   25.640        2   23.000

5   22.000       34   20.250                5   24.800      104   23.000

121   21.900       38   20.200               14   24.600       16   23.000

14   21.800       32   20.200              121   24.400      124   22.900

4   21.500      152   20.200               24   24.400        8   22.700

166   21.400        8   20.100               39   24.300      152   22.700

20   21.330      163   20.000              166   24.300      109   22.640

30   21.100      104   20.000                4   24.200      164   22.620

15   21.100      118   20.000              120   24.000      155   22.580

21   21.090      109   19.990               20   23.960      151   22.500

1   21.000      155   19.980               21   23.920      153   22.410

120   21.000      115   19.900               46   23.830       45   22.310

3   21.000      107   19.900               38   23.800      110   22.300

24   21.000       23   19.880               15   23.800      115   22.200

10   20.900       35   19.770               31   23.800       23   22.100

12   20.900      124   19.500                3   23.600      118   22.000

46   20.880      110   19.500               19   23.600       40   22.000

39   20.800      112   19.100               17   23.500       34   21.850

114   20.800       40   19.000               10   23.500      108   21.800

19   20.800       36   18.800               12   23.500       35   21.670

33   20.700        7   18.760              165   23.500      112   21.400

31   20.630      108   18.600               33   23.300       36   21.300

165   20.600      164   18.500              107   23.200        7   21.300

2   20.500      140   18.100               27   23.100      113   21.000

151   20.500      113   18.000               30   23.100      145   20.100

16   20.500      133   17.530 (*)            1   23.100      140   20.000

17   20.500       22   17.300 (*)          114   23.100       22   18.000 (*)

27   20.500      145   17.100 (*)          158   23.100      133   14.810 (*)

- Figures

Figure 2:	Youden plot of SO2-S in absorbing solution.

Figure 3:	Youden plot of SO2-S on impregnated filter.

Figure 4:	Youden plot of HNO3-N on impregnated filter.

Figure 5:	Youden plot of NO2-N in absorbing solution.

Figure 6:	Youden plot of NH3-N on impregnated filter.

Figure 7:	Youden plot of SO4-S in precipitation.

Figure 8:	Youden plot of NO3-N in precipitation.

Figure 9:	Youden plot of NH4-N in precipitation.

Figure 10:	Youden plot of pH in precipitation.

Figure 11:	Youden plot of strong acid in precipitation.

Figure 12:	Youden plot of Cl in precipitation.

Figure 13:	Youden plot of Na in precipitation.

Figure 14:	Youden plot of Mg in precipitation.

Figure 15:	Youden plot of Ca in precipitation.

Figure 16:	Youden plot of K in precipitation.

Figure 17:	Youden plot of conductivity in precipitation.